## Sample Setup
Block Type | 96-Well 0.2-mL Block
Calibration Background is expired | No
Calibration Background performed on | 06-01-2021
Calibration Pure Dye ABY is expired | No
Calibration Pure Dye ABY performed on | 06-01-2021
Calibration Pure Dye FAM is expired | No
Calibration Pure Dye FAM performed on | 06-01-2021
Calibration Pure Dye JUN is expired | No
Calibration Pure Dye JUN performed on | 06-01-2021
Calibration Pure Dye NED is expired | No
Calibration Pure Dye NED performed on | 06-01-2021
Calibration Pure Dye ROX is expired | No
Calibration Pure Dye ROX performed on | 06-01-2021
Calibration Pure Dye SYBR is expired | No
Calibration Pure Dye SYBR performed on | 06-01-2021
Calibration Pure Dye TAMRA is expired | No
Calibration Pure Dye TAMRA performed on | 06-01-2021
Calibration Pure Dye VIC is expired | No
Calibration Pure Dye VIC performed on | 06-01-2021
Calibration ROI is expired | No
Calibration ROI performed on | 06-01-2021
Calibration Uniformity is expired | No
Calibration Uniformity performed on | 06-01-2021
Chemistry | TAQMAN
Date Created | 2022-03-04 10:21:50 AM CST
Experiment Barcode
Experiment Comment
Experiment File Name | D:\Users\INSTR-ADMIN\Desktop\Zoe QS3 data\1110304\SMA QC test_partA2X_RNP05X_20220304_Zoe.eds
Experiment Name | SMA QC test_partA2X_RNP05X_20220304_Zoe
Experiment Run End Time | 2022-03-04 10:13:25 AM CST
Experiment Type | Standard Curve
Instrument Name | 272325672
Instrument Serial Number | 272325672
Instrument Type | QuantStudio™ 3 System
Passive Reference | ROX
Post-read Stage/Step
Pre-read Stage/Step
Quantification Cycle Method | Ct
Signal Smoothing On | true
Stage/ Cycle where Ct Analysis is performed | Stage2, Step2
User Name
Well | Well Position | Sample Name | Sample Color | Biogroup Name | Biogroup Color | Target Name | Target Color | Task | Reporter | Quencher | Quantity | Comments
1 | A1
2 | A2
3 | A3 | SMN 1-4_gDNA | RGB(0,139,69) |  |  | RNP | RGB(0,139,69) | UNKNOWN | TAMRA | NFQ-MGB
3 | A3 | SMN 1-4_gDNA | RGB(0,139,69) |  |  | SMN1 | RGB(176,23,31) | UNKNOWN | FAM | NFQ-MGB
3 | A3 | SMN 1-4_gDNA | RGB(0,139,69) |  |  | SMN2 | RGB(0,0,255) | UNKNOWN | VIC | NFQ-MGB
4 | A4
5 | A5
6 | A6
7 | A7
8 | A8
9 | A9
10 | A10
11 | A11
12 | A12
13 | B1
14 | B2
15 | B3 | SMN 2-1_gDNA | RGB(255,127,0) |  |  | RNP | RGB(0,139,69) | UNKNOWN | TAMRA | NFQ-MGB
15 | B3 | SMN 2-1_gDNA | RGB(255,127,0) |  |  | SMN1 | RGB(176,23,31) | UNKNOWN | FAM | NFQ-MGB
15 | B3 | SMN 2-1_gDNA | RGB(255,127,0) |  |  | SMN2 | RGB(0,0,255) | UNKNOWN | VIC | NFQ-MGB
16 | B4
17 | B5
18 | B6
19 | B7
20 | B8
21 | B9
22 | B10
23 | B11
24 | B12
25 | C1
26 | C2
27 | C3 | partA-600-12_gDNA | RGB(139,137,112) |  |  | RNP | RGB(0,139,69) | UNKNOWN | TAMRA | NFQ-MGB
27 | C3 | partA-600-12_gDNA | RGB(139,137,112) |  |  | SMN1 | RGB(176,23,31) | UNKNOWN | FAM | NFQ-MGB
27 | C3 | partA-600-12_gDNA | RGB(139,137,112) |  |  | SMN2 | RGB(0,0,255) | UNKNOWN | VIC | NFQ-MGB
28 | C4
29 | C5
30 | C6
31 | C7
32 | C8
33 | C9
34 | C10
35 | C11
36 | C12
37 | D1
38 | D2
39 | D3 | SMN 1+2_gDNA | RGB(0,0,255) |  |  | RNP | RGB(0,139,69) | UNKNOWN | TAMRA | NFQ-MGB
39 | D3 | SMN 1+2_gDNA | RGB(0,0,255) |  |  | SMN1 | RGB(176,23,31) | UNKNOWN | FAM | NFQ-MGB
39 | D3 | SMN 1+2_gDNA | RGB(0,0,255) |  |  | SMN2 | RGB(0,0,255) | UNKNOWN | VIC | NFQ-MGB
40 | D4
41 | D5 | NTC | RGB(176,23,31) |  |  | RNP | RGB(0,139,69) | UNKNOWN | TAMRA | NFQ-MGB
41 | D5 | NTC | RGB(176,23,31) |  |  | SMN1 | RGB(176,23,31) | UNKNOWN | FAM | NFQ-MGB
41 | D5 | NTC | RGB(176,23,31) |  |  | SMN2 | RGB(0,0,255) | UNKNOWN | VIC | NFQ-MGB
42 | D6
43 | D7
44 | D8
45 | D9
46 | D10
47 | D11
48 | D12
49 | E1
50 | E2
51 | E3 | SMN 1+2+partA_gDNA | RGB(176,23,31) |  |  | RNP | RGB(0,139,69) | UNKNOWN | TAMRA | NFQ-MGB
51 | E3 | SMN 1+2+partA_gDNA | RGB(176,23,31) |  |  | SMN1 | RGB(176,23,31) | UNKNOWN | FAM | NFQ-MGB
51 | E3 | SMN 1+2+partA_gDNA | RGB(176,23,31) |  |  | SMN2 | RGB(0,0,255) | UNKNOWN | VIC | NFQ-MGB
52 | E4
53 | E5
54 | E6
55 | E7
56 | E8
57 | E9
58 | E10
59 | E11
60 | E12
61 | F1
62 | F2
63 | F3 | #24 | RGB(176,23,31) |  |  | RNP | RGB(0,139,69) | UNKNOWN | TAMRA | NFQ-MGB
63 | F3 | #24 | RGB(176,23,31) |  |  | SMN1 | RGB(176,23,31) | UNKNOWN | FAM | NFQ-MGB
63 | F3 | #24 | RGB(176,23,31) |  |  | SMN2 | RGB(0,0,255) | UNKNOWN | VIC | NFQ-MGB
64 | F4
65 | F5
66 | F6
67 | F7
68 | F8
69 | F9
70 | F10
71 | F11
72 | F12
73 | G1
74 | G2
75 | G3 | #25 | RGB(0,0,255) |  |  | RNP | RGB(0,139,69) | UNKNOWN | TAMRA | NFQ-MGB
75 | G3 | #25 | RGB(0,0,255) |  |  | SMN1 | RGB(176,23,31) | UNKNOWN | FAM | NFQ-MGB
75 | G3 | #25 | RGB(0,0,255) |  |  | SMN2 | RGB(0,0,255) | UNKNOWN | VIC | NFQ-MGB
76 | G4
77 | G5
78 | G6
79 | G7
80 | G8
81 | G9
82 | G10
83 | G11
84 | G12
85 | H1
86 | H2
87 | H3 | #26 | RGB(0,139,69) |  |  | RNP | RGB(0,139,69) | UNKNOWN | TAMRA | NFQ-MGB
87 | H3 | #26 | RGB(0,139,69) |  |  | SMN1 | RGB(176,23,31) | UNKNOWN | FAM | NFQ-MGB
87 | H3 | #26 | RGB(0,139,69) |  |  | SMN2 | RGB(0,0,255) | UNKNOWN | VIC | NFQ-MGB
88 | H4
89 | H5
90 | H6
91 | H7
92 | H8
93 | H9
94 | H10
95 | H11
96 | H12

## Amplification Data
Block Type | 96-Well 0.2-mL Block
Calibration Background is expired | No
Calibration Background performed on | 06-01-2021
Calibration Pure Dye ABY is expired | No
Calibration Pure Dye ABY performed on | 06-01-2021
Calibration Pure Dye FAM is expired | No
Calibration Pure Dye FAM performed on | 06-01-2021
Calibration Pure Dye JUN is expired | No
Calibration Pure Dye JUN performed on | 06-01-2021
Calibration Pure Dye NED is expired | No
Calibration Pure Dye NED performed on | 06-01-2021
Calibration Pure Dye ROX is expired | No
Calibration Pure Dye ROX performed on | 06-01-2021
Calibration Pure Dye SYBR is expired | No
Calibration Pure Dye SYBR performed on | 06-01-2021
Calibration Pure Dye TAMRA is expired | No
Calibration Pure Dye TAMRA performed on | 06-01-2021
Calibration Pure Dye VIC is expired | No
Calibration Pure Dye VIC performed on | 06-01-2021
Calibration ROI is expired | No
Calibration ROI performed on | 06-01-2021
Calibration Uniformity is expired | No
Calibration Uniformity performed on | 06-01-2021
Chemistry | TAQMAN
Date Created | 2022-03-04 10:21:50 AM CST
Experiment Barcode
Experiment Comment
Experiment File Name | D:\Users\INSTR-ADMIN\Desktop\Zoe QS3 data\1110304\SMA QC test_partA2X_RNP05X_20220304_Zoe.eds
Experiment Name | SMA QC test_partA2X_RNP05X_20220304_Zoe
Experiment Run End Time | 2022-03-04 10:13:25 AM CST
Experiment Type | Standard Curve
Instrument Name | 272325672
Instrument Serial Number | 272325672
Instrument Type | QuantStudio™ 3 System
Passive Reference | ROX
Post-read Stage/Step
Pre-read Stage/Step
Quantification Cycle Method | Ct
Signal Smoothing On | true
Stage/ Cycle where Ct Analysis is performed | Stage2, Step2
User Name
Well | Well Position | Cycle | Target Name | Rn | Delta Rn
1 | A1 | 1
1 | A1 | 2
1 | A1 | 3
1 | A1 | 4
1 | A1 | 5
1 | A1 | 6
1 | A1 | 7
1 | A1 | 8
1 | A1 | 9
1 | A1 | 10
1 | A1 | 11
1 | A1 | 12
1 | A1 | 13
1 | A1 | 14
1 | A1 | 15
1 | A1 | 16
1 | A1 | 17
1 | A1 | 18
1 | A1 | 19
1 | A1 | 20
1 | A1 | 21
1 | A1 | 22
1 | A1 | 23
1 | A1 | 24
1 | A1 | 25
1 | A1 | 26
1 | A1 | 27
1 | A1 | 28
1 | A1 | 29
1 | A1 | 30
1 | A1 | 31
1 | A1 | 32
1 | A1 | 33
1 | A1 | 34
1 | A1 | 35
2 | A2 | 1
2 | A2 | 2
2 | A2 | 3
2 | A2 | 4
2 | A2 | 5
2 | A2 | 6
2 | A2 | 7
2 | A2 | 8
2 | A2 | 9
2 | A2 | 10
2 | A2 | 11
2 | A2 | 12
2 | A2 | 13
2 | A2 | 14
2 | A2 | 15
2 | A2 | 16
2 | A2 | 17
2 | A2 | 18
2 | A2 | 19
2 | A2 | 20
2 | A2 | 21
2 | A2 | 22
2 | A2 | 23
2 | A2 | 24
2 | A2 | 25
2 | A2 | 26
2 | A2 | 27
2 | A2 | 28
2 | A2 | 29
2 | A2 | 30
2 | A2 | 31
2 | A2 | 32
2 | A2 | 33
2 | A2 | 34
2 | A2 | 35
3 | A3 | 1 | RNP | 0.10579 | 0.001666
3 | A3 | 2 | RNP | 0.10584 | 0.0017451
3 | A3 | 3 | RNP | 0.1037 | -0.000356
3 | A3 | 4 | RNP | 0.1033 | -0.00072154
3 | A3 | 5 | RNP | 0.10552 | 0.0015308
3 | A3 | 6 | RNP | 0.10517 | 0.001215
3 | A3 | 7 | RNP | 0.10424 | 0.00032279
3 | A3 | 8 | RNP | 0.10407 | 0.00018151
3 | A3 | 9 | RNP | 0.1041 | 0.00025493
3 | A3 | 10 | RNP | 0.10395 | 0.00013449
3 | A3 | 11 | RNP | 0.10421 | 0.00042883
3 | A3 | 12 | RNP | 0.10221 | -0.0015295
3 | A3 | 13 | RNP | 0.1018 | -0.0019085
3 | A3 | 14 | RNP | 0.1028 | -0.00087886
3 | A3 | 15 | RNP | 0.1033 | -0.00033885
3 | A3 | 16 | RNP | 0.10426 | 0.00065291
3 | A3 | 17 | RNP | 0.10389 | 0.00031898
3 | A3 | 18 | RNP | 0.10173 | -0.0018037
3 | A3 | 19 | RNP | 0.10192 | -0.001578
3 | A3 | 20 | RNP | 0.10496 | 0.001497
3 | A3 | 21 | RNP | 0.10601 | 0.0025778
3 | A3 | 22 | RNP | 0.11001 | 0.0066111
3 | A3 | 23 | RNP | 0.11804 | 0.014683
3 | A3 | 24 | RNP | 0.13046 | 0.027136
3 | A3 | 25 | RNP | 0.14594 | 0.042652
3 | A3 | 26 | RNP | 0.17406 | 0.070809
3 | A3 | 27 | RNP | 0.21523 | 0.11201
3 | A3 | 28 | RNP | 0.26517 | 0.16199
3 | A3 | 29 | RNP | 0.31781 | 0.21466
3 | A3 | 30 | RNP | 0.36815 | 0.26503
3 | A3 | 31 | RNP | 0.4075 | 0.30442
3 | A3 | 32 | RNP | 0.44521 | 0.34216
3 | A3 | 33 | RNP | 0.47962 | 0.3766
3 | A3 | 34 | RNP | 0.50658 | 0.4036
3 | A3 | 35 | RNP | 0.52512 | 0.42218
3 | A3 | 1 | SMN2 | 0.34016 | 0.016101
3 | A3 | 2 | SMN2 | 0.33861 | 0.013182
3 | A3 | 3 | SMN2 | 0.33721 | 0.01041
3 | A3 | 4 | SMN2 | 0.33834 | 0.010169
3 | A3 | 5 | SMN2 | 0.33757 | 0.0080299
3 | A3 | 6 | SMN2 | 0.33708 | 0.0061695
3 | A3 | 7 | SMN2 | 0.33756 | 0.0052758
3 | A3 | 8 | SMN2 | 0.34075 | 0.0071033
3 | A3 | 9 | SMN2 | 0.34145 | 0.0064262
3 | A3 | 10 | SMN2 | 0.3416 | 0.0052102
3 | A3 | 11 | SMN2 | 0.341 | 0.0032442
3 | A3 | 12 | SMN2 | 0.33986 | 0.00073361
3 | A3 | 13 | SMN2 | 0.33617 | -0.0043299
3 | A3 | 14 | SMN2 | 0.33555 | -0.0063198
3 | A3 | 15 | SMN2 | 0.33689 | -0.0063485
3 | A3 | 16 | SMN2 | 0.33731 | -0.0072977
3 | A3 | 17 | SMN2 | 0.33794 | -0.0080364
3 | A3 | 18 | SMN2 | 0.33754 | -0.009809
3 | A3 | 19 | SMN2 | 0.3356 | -0.013118
3 | A3 | 20 | SMN2 | 0.33619 | -0.013896
3 | A3 | 21 | SMN2 | 0.33973 | -0.011729
3 | A3 | 22 | SMN2 | 0.34116 | -0.011665
3 | A3 | 23 | SMN2 | 0.34232 | -0.011872
3 | A3 | 24 | SMN2 | 0.34278 | -0.012789
3 | A3 | 25 | SMN2 | 0.34745 | -0.009482
3 | A3 | 26 | SMN2 | 0.35214 | -0.0061627
3 | A3 | 27 | SMN2 | 0.35631 | -0.003363
3 | A3 | 28 | SMN2 | 0.35981 | -0.001236
3 | A3 | 29 | SMN2 | 0.36516 | 0.0027441
3 | A3 | 30 | SMN2 | 0.37238 | 0.008599
3 | A3 | 31 | SMN2 | 0.37824 | 0.013084
3 | A3 | 32 | SMN2 | 0.38111 | 0.014583
3 | A3 | 33 | SMN2 | 0.38511 | 0.017213
3 | A3 | 34 | SMN2 | 0.38772 | 0.018459
3 | A3 | 35 | SMN2 | 0.39265 | 0.022015
3 | A3 | 1 | SMN1 | 0.83655 | -0.013652
3 | A3 | 2 | SMN1 | 0.83641 | -0.013026
3 | A3 | 3 | SMN1 | 0.84231 | -0.0063617
3 | A3 | 4 | SMN1 | 0.84787 | -2.9744e-05
3 | A3 | 5 | SMN1 | 0.84931 | 0.0021746
3 | A3 | 6 | SMN1 | 0.84634 | -2.9662e-05
3 | A3 | 7 | SMN1 | 0.84655 | 0.00093898
3 | A3 | 8 | SMN1 | 0.84964 | 0.0047997
3 | A3 | 9 | SMN1 | 0.84952 | 0.0054467
3 | A3 | 10 | SMN1 | 0.84781 | 0.0045036
3 | A3 | 11 | SMN1 | 0.84486 | 0.0023113
3 | A3 | 12 | SMN1 | 0.84073 | -0.0010464
3 | A3 | 13 | SMN1 | 0.83533 | -0.0056884
3 | A3 | 14 | SMN1 | 0.83424 | -0.0060117
3 | A3 | 15 | SMN1 | 0.83375 | -0.0057333
3 | A3 | 16 | SMN1 | 0.83454 | -0.0041787
3 | A3 | 17 | SMN1 | 0.83724 | -0.00071344
3 | A3 | 18 | SMN1 | 0.84022 | 0.0030311
3 | A3 | 19 | SMN1 | 0.84301 | 0.0065872
3 | A3 | 20 | SMN1 | 0.84651 | 0.010848
3 | A3 | 21 | SMN1 | 0.85569 | 0.020794
3 | A3 | 22 | SMN1 | 0.87991 | 0.04578
3 | A3 | 23 | SMN1 | 0.92378 | 0.09042
3 | A3 | 24 | SMN1 | 0.99945 | 0.16685
3 | A3 | 25 | SMN1 | 1.1331 | 0.30131
3 | A3 | 26 | SMN1 | 1.3124 | 0.4813
3 | A3 | 27 | SMN1 | 1.5247 | 0.69442
3 | A3 | 28 | SMN1 | 1.74 | 0.91043
3 | A3 | 29 | SMN1 | 1.9672 | 1.1384
3 | A3 | 30 | SMN1 | 2.1752 | 1.3472
3 | A3 | 31 | SMN1 | 2.3805 | 1.5533
3 | A3 | 32 | SMN1 | 2.5853 | 1.7588
3 | A3 | 33 | SMN1 | 2.7903 | 1.9646
3 | A3 | 34 | SMN1 | 2.992 | 2.1671
3 | A3 | 35 | SMN1 | 3.1552 | 2.331
4 | A4 | 1
4 | A4 | 2
4 | A4 | 3
4 | A4 | 4
4 | A4 | 5
4 | A4 | 6
4 | A4 | 7
4 | A4 | 8
4 | A4 | 9
4 | A4 | 10
4 | A4 | 11
4 | A4 | 12
4 | A4 | 13
4 | A4 | 14
4 | A4 | 15
4 | A4 | 16
4 | A4 | 17
4 | A4 | 18
4 | A4 | 19
4 | A4 | 20
4 | A4 | 21
4 | A4 | 22
4 | A4 | 23
4 | A4 | 24
4 | A4 | 25
4 | A4 | 26
4 | A4 | 27
4 | A4 | 28
4 | A4 | 29
4 | A4 | 30
4 | A4 | 31
4 | A4 | 32
4 | A4 | 33
4 | A4 | 34
4 | A4 | 35
5 | A5 | 1
5 | A5 | 2
5 | A5 | 3
5 | A5 | 4
5 | A5 | 5
5 | A5 | 6
5 | A5 | 7
5 | A5 | 8
5 | A5 | 9
5 | A5 | 10
5 | A5 | 11
5 | A5 | 12
5 | A5 | 13
5 | A5 | 14
5 | A5 | 15
5 | A5 | 16
5 | A5 | 17
5 | A5 | 18
5 | A5 | 19
5 | A5 | 20
5 | A5 | 21
5 | A5 | 22
5 | A5 | 23
5 | A5 | 24
5 | A5 | 25
5 | A5 | 26
5 | A5 | 27
5 | A5 | 28
5 | A5 | 29
5 | A5 | 30
5 | A5 | 31
5 | A5 | 32
5 | A5 | 33
5 | A5 | 34
5 | A5 | 35
6 | A6 | 1
6 | A6 | 2
6 | A6 | 3
6 | A6 | 4
6 | A6 | 5
6 | A6 | 6
6 | A6 | 7
6 | A6 | 8
6 | A6 | 9
6 | A6 | 10
6 | A6 | 11
6 | A6 | 12
6 | A6 | 13
6 | A6 | 14
6 | A6 | 15
6 | A6 | 16
6 | A6 | 17
6 | A6 | 18
6 | A6 | 19
6 | A6 | 20
6 | A6 | 21
6 | A6 | 22
6 | A6 | 23
6 | A6 | 24
6 | A6 | 25
6 | A6 | 26
6 | A6 | 27
6 | A6 | 28
6 | A6 | 29
6 | A6 | 30
6 | A6 | 31
6 | A6 | 32
6 | A6 | 33
6 | A6 | 34
6 | A6 | 35
7 | A7 | 1
7 | A7 | 2
7 | A7 | 3
7 | A7 | 4
7 | A7 | 5
7 | A7 | 6
7 | A7 | 7
7 | A7 | 8
7 | A7 | 9
7 | A7 | 10
7 | A7 | 11
7 | A7 | 12
7 | A7 | 13
7 | A7 | 14
7 | A7 | 15
7 | A7 | 16
7 | A7 | 17
7 | A7 | 18
7 | A7 | 19
7 | A7 | 20
7 | A7 | 21
7 | A7 | 22
7 | A7 | 23
7 | A7 | 24
7 | A7 | 25
7 | A7 | 26
7 | A7 | 27
7 | A7 | 28
7 | A7 | 29
7 | A7 | 30
7 | A7 | 31
7 | A7 | 32
7 | A7 | 33
7 | A7 | 34
7 | A7 | 35
8 | A8 | 1
8 | A8 | 2
8 | A8 | 3
8 | A8 | 4
8 | A8 | 5
8 | A8 | 6
8 | A8 | 7
8 | A8 | 8
8 | A8 | 9
8 | A8 | 10
8 | A8 | 11
8 | A8 | 12
8 | A8 | 13
8 | A8 | 14
8 | A8 | 15
8 | A8 | 16
8 | A8 | 17
8 | A8 | 18
8 | A8 | 19
8 | A8 | 20
8 | A8 | 21
8 | A8 | 22
8 | A8 | 23
8 | A8 | 24
8 | A8 | 25
8 | A8 | 26
8 | A8 | 27
8 | A8 | 28
8 | A8 | 29
8 | A8 | 30
8 | A8 | 31
8 | A8 | 32
8 | A8 | 33
8 | A8 | 34
8 | A8 | 35
9 | A9 | 1
9 | A9 | 2
9 | A9 | 3
9 | A9 | 4
9 | A9 | 5
9 | A9 | 6
9 | A9 | 7
9 | A9 | 8
9 | A9 | 9
9 | A9 | 10
9 | A9 | 11
9 | A9 | 12
9 | A9 | 13
9 | A9 | 14
9 | A9 | 15
9 | A9 | 16
9 | A9 | 17
9 | A9 | 18
9 | A9 | 19
9 | A9 | 20
9 | A9 | 21
9 | A9 | 22
9 | A9 | 23
9 | A9 | 24
9 | A9 | 25
9 | A9 | 26
9 | A9 | 27
9 | A9 | 28
9 | A9 | 29
9 | A9 | 30
9 | A9 | 31
9 | A9 | 32
9 | A9 | 33
9 | A9 | 34
9 | A9 | 35
10 | A10 | 1
10 | A10 | 2
10 | A10 | 3
10 | A10 | 4
10 | A10 | 5
10 | A10 | 6
10 | A10 | 7
10 | A10 | 8
10 | A10 | 9
10 | A10 | 10
10 | A10 | 11
10 | A10 | 12
10 | A10 | 13
10 | A10 | 14
10 | A10 | 15
10 | A10 | 16
10 | A10 | 17
10 | A10 | 18
10 | A10 | 19
10 | A10 | 20
10 | A10 | 21
10 | A10 | 22
10 | A10 | 23
10 | A10 | 24
10 | A10 | 25
10 | A10 | 26
10 | A10 | 27
10 | A10 | 28
10 | A10 | 29
10 | A10 | 30
10 | A10 | 31
10 | A10 | 32
10 | A10 | 33
10 | A10 | 34
10 | A10 | 35
11 | A11 | 1
11 | A11 | 2
11 | A11 | 3
11 | A11 | 4
11 | A11 | 5
11 | A11 | 6
11 | A11 | 7
11 | A11 | 8
11 | A11 | 9
11 | A11 | 10
11 | A11 | 11
11 | A11 | 12
11 | A11 | 13
11 | A11 | 14
11 | A11 | 15
11 | A11 | 16
11 | A11 | 17
11 | A11 | 18
11 | A11 | 19
11 | A11 | 20
11 | A11 | 21
11 | A11 | 22
11 | A11 | 23
11 | A11 | 24
11 | A11 | 25
11 | A11 | 26
11 | A11 | 27
11 | A11 | 28
11 | A11 | 29
11 | A11 | 30
11 | A11 | 31
11 | A11 | 32
11 | A11 | 33
11 | A11 | 34
11 | A11 | 35
12 | A12 | 1
12 | A12 | 2
12 | A12 | 3
12 | A12 | 4
12 | A12 | 5
12 | A12 | 6
12 | A12 | 7
12 | A12 | 8
12 | A12 | 9
12 | A12 | 10
12 | A12 | 11
12 | A12 | 12
12 | A12 | 13
12 | A12 | 14
12 | A12 | 15
12 | A12 | 16
12 | A12 | 17
12 | A12 | 18
12 | A12 | 19
12 | A12 | 20
12 | A12 | 21
12 | A12 | 22
12 | A12 | 23
12 | A12 | 24
12 | A12 | 25
12 | A12 | 26
12 | A12 | 27
12 | A12 | 28
12 | A12 | 29
12 | A12 | 30
12 | A12 | 31
12 | A12 | 32
12 | A12 | 33
12 | A12 | 34
12 | A12 | 35
13 | B1 | 1
13 | B1 | 2
13 | B1 | 3
13 | B1 | 4
13 | B1 | 5
13 | B1 | 6
13 | B1 | 7
13 | B1 | 8
13 | B1 | 9
13 | B1 | 10
13 | B1 | 11
13 | B1 | 12
13 | B1 | 13
13 | B1 | 14
13 | B1 | 15
13 | B1 | 16
13 | B1 | 17
13 | B1 | 18
13 | B1 | 19
13 | B1 | 20
13 | B1 | 21
13 | B1 | 22
13 | B1 | 23
13 | B1 | 24
13 | B1 | 25
13 | B1 | 26
13 | B1 | 27
13 | B1 | 28
13 | B1 | 29
13 | B1 | 30
13 | B1 | 31
13 | B1 | 32
13 | B1 | 33
13 | B1 | 34
13 | B1 | 35
14 | B2 | 1
14 | B2 | 2
14 | B2 | 3
14 | B2 | 4
14 | B2 | 5
14 | B2 | 6
14 | B2 | 7
14 | B2 | 8
14 | B2 | 9
14 | B2 | 10
14 | B2 | 11
14 | B2 | 12
14 | B2 | 13
14 | B2 | 14
14 | B2 | 15
14 | B2 | 16
14 | B2 | 17
14 | B2 | 18
14 | B2 | 19
14 | B2 | 20
14 | B2 | 21
14 | B2 | 22
14 | B2 | 23
14 | B2 | 24
14 | B2 | 25
14 | B2 | 26
14 | B2 | 27
14 | B2 | 28
14 | B2 | 29
14 | B2 | 30
14 | B2 | 31
14 | B2 | 32
14 | B2 | 33
14 | B2 | 34
14 | B2 | 35
15 | B3 | 1 | RNP | 0.090248 | 0.00034522
15 | B3 | 2 | RNP | 0.090915 | 0.00075354
15 | B3 | 3 | RNP | 0.09132 | 0.00089996
15 | B3 | 4 | RNP | 0.092674 | 0.0019962
15 | B3 | 5 | RNP | 0.092573 | 0.0016369
15 | B3 | 6 | RNP | 0.091993 | 0.00079778
15 | B3 | 7 | RNP | 0.091905 | 0.00045144
15 | B3 | 8 | RNP | 0.092982 | 0.0012703
15 | B3 | 9 | RNP | 0.092238 | 0.00026818
15 | B3 | 10 | RNP | 0.091722 | -0.00050643
15 | B3 | 11 | RNP | 0.091142 | -0.0013449
15 | B3 | 12 | RNP | 0.090572 | -0.0021728
15 | B3 | 13 | RNP | 0.090696 | -0.0023072
15 | B3 | 14 | RNP | 0.090742 | -0.00252
15 | B3 | 15 | RNP | 0.09112 | -0.0024002
15 | B3 | 16 | RNP | 0.091864 | -0.0019146
15 | B3 | 17 | RNP | 0.092292 | -0.0017455
15 | B3 | 18 | RNP | 0.093867 | -0.0004291
15 | B3 | 19 | RNP | 0.095 | 0.00044549
15 | B3 | 20 | RNP | 0.095449 | 0.00063636
15 | B3 | 21 | RNP | 0.096632 | 0.0015607
15 | B3 | 22 | RNP | 0.097129 | 0.0017995
15 | B3 | 23 | RNP | 0.099166 | 0.0035779
15 | B3 | 24 | RNP | 0.1058 | 0.0099563
15 | B3 | 25 | RNP | 0.11783 | 0.021724
15 | B3 | 26 | RNP | 0.1402 | 0.043834
15 | B3 | 27 | RNP | 0.1746 | 0.077977
15 | B3 | 28 | RNP | 0.21963 | 0.12275
15 | B3 | 29 | RNP | 0.26962 | 0.17248
15 | B3 | 30 | RNP | 0.31812 | 0.22072
15 | B3 | 31 | RNP | 0.36377 | 0.26612
15 | B3 | 32 | RNP | 0.40431 | 0.3064
15 | B3 | 33 | RNP | 0.44227 | 0.3441
15 | B3 | 34 | RNP | 0.47185 | 0.37342
15 | B3 | 35 | RNP | 0.49216 | 0.39347
15 | B3 | 1 | SMN2 | 0.33301 | 0.0025589
15 | B3 | 2 | SMN2 | 0.33149 | 0.00095771
15 | B3 | 3 | SMN2 | 0.32959 | -0.0010237
15 | B3 | 4 | SMN2 | 0.32975 | -0.00093876
15 | B3 | 5 | SMN2 | 0.33207 | 0.0012994
15 | B3 | 6 | SMN2 | 0.33241 | 0.0015563
15 | B3 | 7 | SMN2 | 0.33057 | -0.00036217
15 | B3 | 8 | SMN2 | 0.33016 | -0.00084774
15 | B3 | 9 | SMN2 | 0.33081 | -0.00027765
15 | B3 | 10 | SMN2 | 0.33107 | -9.9046e-05
15 | B3 | 11 | SMN2 | 0.33136 | 0.00011541
15 | B3 | 12 | SMN2 | 0.33179 | 0.00046562
15 | B3 | 13 | SMN2 | 0.33323 | 0.0018214
15 | B3 | 14 | SMN2 | 0.3329 | 0.0014097
15 | B3 | 15 | SMN2 | 0.33057 | -0.0010001
15 | B3 | 16 | SMN2 | 0.33025 | -0.0013969
15 | B3 | 17 | SMN2 | 0.33052 | -0.0012066
15 | B3 | 18 | SMN2 | 0.33091 | -0.00089503
15 | B3 | 19 | SMN2 | 0.33198 | 9.5182e-05
15 | B3 | 20 | SMN2 | 0.33325 | 0.0012846
15 | B3 | 21 | SMN2 | 0.3375 | 0.0054496
15 | B3 | 22 | SMN2 | 0.3451 | 0.012977
15 | B3 | 23 | SMN2 | 0.3553 | 0.023097
15 | B3 | 24 | SMN2 | 0.3756 | 0.04331
15 | B3 | 25 | SMN2 | 0.41982 | 0.087448
15 | B3 | 26 | SMN2 | 0.49967 | 0.16722
15 | B3 | 27 | SMN2 | 0.62606 | 0.29353
15 | B3 | 28 | SMN2 | 0.77753 | 0.44492
15 | B3 | 29 | SMN2 | 0.94089 | 0.6082
15 | B3 | 30 | SMN2 | 1.0937 | 0.76095
15 | B3 | 31 | SMN2 | 1.2346 | 0.90179
15 | B3 | 32 | SMN2 | 1.3612 | 1.0282
15 | B3 | 33 | SMN2 | 1.4829 | 1.1499
15 | B3 | 34 | SMN2 | 1.5896 | 1.2565
15 | B3 | 35 | SMN2 | 1.6691 | 1.3359
15 | B3 | 1 | SMN1 | 0.48761 | 0.0016088
15 | B3 | 2 | SMN1 | 0.48841 | 0.0015988
15 | B3 | 3 | SMN1 | 0.49025 | 0.0026271
15 | B3 | 4 | SMN1 | 0.49531 | 0.0068774
15 | B3 | 5 | SMN1 | 0.49512 | 0.0058722
15 | B3 | 6 | SMN1 | 0.49461 | 0.0045575
15 | B3 | 7 | SMN1 | 0.49352 | 0.0026513
15 | B3 | 8 | SMN1 | 0.48983 | -0.0018447
15 | B3 | 9 | SMN1 | 0.48912 | -0.0033703
15 | B3 | 10 | SMN1 | 0.49057 | -0.0027284
15 | B3 | 11 | SMN1 | 0.49207 | -0.002037
15 | B3 | 12 | SMN1 | 0.49396 | -0.00095997
15 | B3 | 13 | SMN1 | 0.49421 | -0.0015124
15 | B3 | 14 | SMN1 | 0.49143 | -0.0051061
15 | B3 | 15 | SMN1 | 0.49265 | -0.0046932
15 | B3 | 16 | SMN1 | 0.4966 | -0.0015571
15 | B3 | 17 | SMN1 | 0.49908 | 0.00011314
15 | B3 | 18 | SMN1 | 0.49819 | -0.0015893
15 | B3 | 19 | SMN1 | 0.49861 | -0.0019757
15 | B3 | 20 | SMN1 | 0.499 | -0.0023935
15 | B3 | 21 | SMN1 | 0.49993 | -0.0022786
15 | B3 | 22 | SMN1 | 0.5032 | 0.00017937
15 | B3 | 23 | SMN1 | 0.50358 | -0.00025042
15 | B3 | 24 | SMN1 | 0.50398 | -0.00065566
15 | B3 | 25 | SMN1 | 0.50543 | -1.4716e-05
15 | B3 | 26 | SMN1 | 0.50645 | 0.00018944
15 | B3 | 27 | SMN1 | 0.5084 | 0.0013339
15 | B3 | 28 | SMN1 | 0.50476 | -0.0031196
15 | B3 | 29 | SMN1 | 0.50469 | -0.0040015
15 | B3 | 30 | SMN1 | 0.50925 | -0.0002527
15 | B3 | 31 | SMN1 | 0.51066 | 0.00035445
15 | B3 | 32 | SMN1 | 0.5138 | 0.0026844
15 | B3 | 33 | SMN1 | 0.51745 | 0.0055167
15 | B3 | 34 | SMN1 | 0.52012 | 0.007384
15 | B3 | 35 | SMN1 | 0.5218 | 0.0082484
16 | B4 | 1
16 | B4 | 2
16 | B4 | 3
16 | B4 | 4
16 | B4 | 5
16 | B4 | 6
16 | B4 | 7
16 | B4 | 8
16 | B4 | 9
16 | B4 | 10
16 | B4 | 11
16 | B4 | 12
16 | B4 | 13
16 | B4 | 14
16 | B4 | 15
16 | B4 | 16
16 | B4 | 17
16 | B4 | 18
16 | B4 | 19
16 | B4 | 20
16 | B4 | 21
16 | B4 | 22
16 | B4 | 23
16 | B4 | 24
16 | B4 | 25
16 | B4 | 26
16 | B4 | 27
16 | B4 | 28
16 | B4 | 29
16 | B4 | 30
16 | B4 | 31
16 | B4 | 32
16 | B4 | 33
16 | B4 | 34
16 | B4 | 35
17 | B5 | 1
17 | B5 | 2
17 | B5 | 3
17 | B5 | 4
17 | B5 | 5
17 | B5 | 6
17 | B5 | 7
17 | B5 | 8
17 | B5 | 9
17 | B5 | 10
17 | B5 | 11
17 | B5 | 12
17 | B5 | 13
17 | B5 | 14
17 | B5 | 15
17 | B5 | 16
17 | B5 | 17
17 | B5 | 18
17 | B5 | 19
17 | B5 | 20
17 | B5 | 21
17 | B5 | 22
17 | B5 | 23
17 | B5 | 24
17 | B5 | 25
17 | B5 | 26
17 | B5 | 27
17 | B5 | 28
17 | B5 | 29
17 | B5 | 30
17 | B5 | 31
17 | B5 | 32
17 | B5 | 33
17 | B5 | 34
17 | B5 | 35
18 | B6 | 1
18 | B6 | 2
18 | B6 | 3
18 | B6 | 4
18 | B6 | 5
18 | B6 | 6
18 | B6 | 7
18 | B6 | 8
18 | B6 | 9
18 | B6 | 10
18 | B6 | 11
18 | B6 | 12
18 | B6 | 13
18 | B6 | 14
18 | B6 | 15
18 | B6 | 16
18 | B6 | 17
18 | B6 | 18
18 | B6 | 19
18 | B6 | 20
18 | B6 | 21
18 | B6 | 22
18 | B6 | 23
18 | B6 | 24
18 | B6 | 25
18 | B6 | 26
18 | B6 | 27
18 | B6 | 28
18 | B6 | 29
18 | B6 | 30
18 | B6 | 31
18 | B6 | 32
18 | B6 | 33
18 | B6 | 34
18 | B6 | 35
19 | B7 | 1
19 | B7 | 2
19 | B7 | 3
19 | B7 | 4
19 | B7 | 5
19 | B7 | 6
19 | B7 | 7
19 | B7 | 8
19 | B7 | 9
19 | B7 | 10
19 | B7 | 11
19 | B7 | 12
19 | B7 | 13
19 | B7 | 14
19 | B7 | 15
19 | B7 | 16
19 | B7 | 17
19 | B7 | 18
19 | B7 | 19
19 | B7 | 20
19 | B7 | 21
19 | B7 | 22
19 | B7 | 23
19 | B7 | 24
19 | B7 | 25
19 | B7 | 26
19 | B7 | 27
19 | B7 | 28
19 | B7 | 29
19 | B7 | 30
19 | B7 | 31
19 | B7 | 32
19 | B7 | 33
19 | B7 | 34
19 | B7 | 35
20 | B8 | 1
20 | B8 | 2
20 | B8 | 3
20 | B8 | 4
20 | B8 | 5
20 | B8 | 6
20 | B8 | 7
20 | B8 | 8
20 | B8 | 9
20 | B8 | 10
20 | B8 | 11
20 | B8 | 12
20 | B8 | 13
20 | B8 | 14
20 | B8 | 15
20 | B8 | 16
20 | B8 | 17
20 | B8 | 18
20 | B8 | 19
20 | B8 | 20
20 | B8 | 21
20 | B8 | 22
20 | B8 | 23
20 | B8 | 24
20 | B8 | 25
20 | B8 | 26
20 | B8 | 27
20 | B8 | 28
20 | B8 | 29
20 | B8 | 30
20 | B8 | 31
20 | B8 | 32
20 | B8 | 33
20 | B8 | 34
20 | B8 | 35
21 | B9 | 1
21 | B9 | 2
21 | B9 | 3
21 | B9 | 4
21 | B9 | 5
21 | B9 | 6
21 | B9 | 7
21 | B9 | 8
21 | B9 | 9
21 | B9 | 10
21 | B9 | 11
21 | B9 | 12
21 | B9 | 13
21 | B9 | 14
21 | B9 | 15
21 | B9 | 16
21 | B9 | 17
21 | B9 | 18
21 | B9 | 19
21 | B9 | 20
21 | B9 | 21
21 | B9 | 22
21 | B9 | 23
21 | B9 | 24
21 | B9 | 25
21 | B9 | 26
21 | B9 | 27
21 | B9 | 28
21 | B9 | 29
21 | B9 | 30
21 | B9 | 31
21 | B9 | 32
21 | B9 | 33
21 | B9 | 34
21 | B9 | 35
22 | B10 | 1
22 | B10 | 2
22 | B10 | 3
22 | B10 | 4
22 | B10 | 5
22 | B10 | 6
22 | B10 | 7
22 | B10 | 8
22 | B10 | 9
22 | B10 | 10
22 | B10 | 11
22 | B10 | 12
22 | B10 | 13
22 | B10 | 14
22 | B10 | 15
22 | B10 | 16
22 | B10 | 17
22 | B10 | 18
22 | B10 | 19
22 | B10 | 20
22 | B10 | 21
22 | B10 | 22
22 | B10 | 23
22 | B10 | 24
22 | B10 | 25
22 | B10 | 26
22 | B10 | 27
22 | B10 | 28
22 | B10 | 29
22 | B10 | 30
22 | B10 | 31
22 | B10 | 32
22 | B10 | 33
22 | B10 | 34
22 | B10 | 35
23 | B11 | 1
23 | B11 | 2
23 | B11 | 3
23 | B11 | 4
23 | B11 | 5
23 | B11 | 6
23 | B11 | 7
23 | B11 | 8
23 | B11 | 9
23 | B11 | 10
23 | B11 | 11
23 | B11 | 12
23 | B11 | 13
23 | B11 | 14
23 | B11 | 15
23 | B11 | 16
23 | B11 | 17
23 | B11 | 18
23 | B11 | 19
23 | B11 | 20
23 | B11 | 21
23 | B11 | 22
23 | B11 | 23
23 | B11 | 24
23 | B11 | 25
23 | B11 | 26
23 | B11 | 27
23 | B11 | 28
23 | B11 | 29
23 | B11 | 30
23 | B11 | 31
23 | B11 | 32
23 | B11 | 33
23 | B11 | 34
23 | B11 | 35
24 | B12 | 1
24 | B12 | 2
24 | B12 | 3
24 | B12 | 4
24 | B12 | 5
24 | B12 | 6
24 | B12 | 7
24 | B12 | 8
24 | B12 | 9
24 | B12 | 10
24 | B12 | 11
24 | B12 | 12
24 | B12 | 13
24 | B12 | 14
24 | B12 | 15
24 | B12 | 16
24 | B12 | 17
24 | B12 | 18
24 | B12 | 19
24 | B12 | 20
24 | B12 | 21
24 | B12 | 22
24 | B12 | 23
24 | B12 | 24
24 | B12 | 25
24 | B12 | 26
24 | B12 | 27
24 | B12 | 28
24 | B12 | 29
24 | B12 | 30
24 | B12 | 31
24 | B12 | 32
24 | B12 | 33
24 | B12 | 34
24 | B12 | 35
25 | C1 | 1
25 | C1 | 2
25 | C1 | 3
25 | C1 | 4
25 | C1 | 5
25 | C1 | 6
25 | C1 | 7
25 | C1 | 8
25 | C1 | 9
25 | C1 | 10
25 | C1 | 11
25 | C1 | 12
25 | C1 | 13
25 | C1 | 14
25 | C1 | 15
25 | C1 | 16
25 | C1 | 17
25 | C1 | 18
25 | C1 | 19
25 | C1 | 20
25 | C1 | 21
25 | C1 | 22
25 | C1 | 23
25 | C1 | 24
25 | C1 | 25
25 | C1 | 26
25 | C1 | 27
25 | C1 | 28
25 | C1 | 29
25 | C1 | 30
25 | C1 | 31
25 | C1 | 32
25 | C1 | 33
25 | C1 | 34
25 | C1 | 35
26 | C2 | 1
26 | C2 | 2
26 | C2 | 3
26 | C2 | 4
26 | C2 | 5
26 | C2 | 6
26 | C2 | 7
26 | C2 | 8
26 | C2 | 9
26 | C2 | 10
26 | C2 | 11
26 | C2 | 12
26 | C2 | 13
26 | C2 | 14
26 | C2 | 15
26 | C2 | 16
26 | C2 | 17
26 | C2 | 18
26 | C2 | 19
26 | C2 | 20
26 | C2 | 21
26 | C2 | 22
26 | C2 | 23
26 | C2 | 24
26 | C2 | 25
26 | C2 | 26
26 | C2 | 27
26 | C2 | 28
26 | C2 | 29
26 | C2 | 30
26 | C2 | 31
26 | C2 | 32
26 | C2 | 33
26 | C2 | 34
26 | C2 | 35
27 | C3 | 1 | RNP | 0.090211 | -0.0058514
27 | C3 | 2 | RNP | 0.090312 | -0.005237
27 | C3 | 3 | RNP | 0.091217 | -0.0038175
27 | C3 | 4 | RNP | 0.090664 | -0.0038568
27 | C3 | 5 | RNP | 0.08991 | -0.0040974
27 | C3 | 6 | RNP | 0.091375 | -0.0021181
27 | C3 | 7 | RNP | 0.091274 | -0.0017055
27 | C3 | 8 | RNP | 0.092084 | -0.00038158
27 | C3 | 9 | RNP | 0.091897 | -5.4789e-05
27 | C3 | 10 | RNP | 0.09203 | 0.0005921
27 | C3 | 11 | RNP | 0.09205 | 0.0011258
27 | C3 | 12 | RNP | 0.091284 | 0.00087358
27 | C3 | 13 | RNP | 0.090675 | 0.00077922
27 | C3 | 14 | RNP | 0.090641 | 0.001259
27 | C3 | 15 | RNP | 0.090399 | 0.0015312
27 | C3 | 16 | RNP | 0.090336 | 0.0019814
27 | C3 | 17 | RNP | 0.091071 | 0.0032309
27 | C3 | 18 | RNP | 0.090638 | 0.0033116
27 | C3 | 19 | RNP | 0.089927 | 0.0031145
27 | C3 | 20 | RNP | 0.089461 | 0.0031618
27 | C3 | 21 | RNP | 0.089778 | 0.0039931
27 | C3 | 22 | RNP | 0.088514 | 0.0032425
27 | C3 | 23 | RNP | 0.087329 | 0.0025714
27 | C3 | 24 | RNP | 0.086921 | 0.0026782
27 | C3 | 25 | RNP | 0.086889 | 0.0031598
27 | C3 | 26 | RNP | 0.083382 | 0.00016652
27 | C3 | 27 | RNP | 0.081763 | -0.00093868
27 | C3 | 28 | RNP | 0.080993 | -0.0011949
27 | C3 | 29 | RNP | 0.078721 | -0.0029526
27 | C3 | 30 | RNP | 0.077447 | -0.0037127
27 | C3 | 31 | RNP | 0.075521 | -0.0051249
27 | C3 | 32 | RNP | 0.075485 | -0.0046469
27 | C3 | 33 | RNP | 0.076984 | -0.0026341
27 | C3 | 34 | RNP | 0.079568 | 0.00046399
27 | C3 | 35 | RNP | 0.081213 | 0.0026229
27 | C3 | 1 | SMN2 | 0.32758 | 0.0030626
27 | C3 | 2 | SMN2 | 0.32702 | 0.0022532
27 | C3 | 3 | SMN2 | 0.32596 | 0.00094782
27 | C3 | 4 | SMN2 | 0.32645 | 0.0011973
27 | C3 | 5 | SMN2 | 0.32681 | 0.0013172
27 | C3 | 6 | SMN2 | 0.32766 | 0.0019163
27 | C3 | 7 | SMN2 | 0.32616 | 0.00017736
27 | C3 | 8 | SMN2 | 0.32543 | -0.0008
27 | C3 | 9 | SMN2 | 0.32528 | -0.001187
27 | C3 | 10 | SMN2 | 0.32527 | -0.0014409
27 | C3 | 11 | SMN2 | 0.32575 | -0.0012123
27 | C3 | 12 | SMN2 | 0.32604 | -0.0011649
27 | C3 | 13 | SMN2 | 0.3259 | -0.00154
27 | C3 | 14 | SMN2 | 0.32678 | -0.00090396
27 | C3 | 15 | SMN2 | 0.32733 | -0.0005992
27 | C3 | 16 | SMN2 | 0.3278 | -0.00037851
27 | C3 | 17 | SMN2 | 0.32633 | -0.0020906
27 | C3 | 18 | SMN2 | 0.3275 | -0.0011617
27 | C3 | 19 | SMN2 | 0.33135 | 0.0024446
27 | C3 | 20 | SMN2 | 0.33363 | 0.0044786
27 | C3 | 21 | SMN2 | 0.33666 | 0.0072625
27 | C3 | 22 | SMN2 | 0.34587 | 0.016238
27 | C3 | 23 | SMN2 | 0.36609 | 0.036208
27 | C3 | 24 | SMN2 | 0.40489 | 0.07477
27 | C3 | 25 | SMN2 | 0.47055 | 0.14019
27 | C3 | 26 | SMN2 | 0.56574 | 0.23513
27 | C3 | 27 | SMN2 | 0.68189 | 0.35103
27 | C3 | 28 | SMN2 | 0.79835 | 0.46725
27 | C3 | 29 | SMN2 | 0.91493 | 0.58359
27 | C3 | 30 | SMN2 | 1.0276 | 0.69599
27 | C3 | 31 | SMN2 | 1.1263 | 0.79449
27 | C3 | 32 | SMN2 | 1.2227 | 0.89064
27 | C3 | 33 | SMN2 | 1.3091 | 0.9768
27 | C3 | 34 | SMN2 | 1.3825 | 1.0499
27 | C3 | 35 | SMN2 | 1.4408 | 1.108
27 | C3 | 1 | SMN1 | 0.48268 | 0.0054556
27 | C3 | 2 | SMN1 | 0.48175 | 0.0039792
27 | C3 | 3 | SMN1 | 0.48055 | 0.0022215
27 | C3 | 4 | SMN1 | 0.4814 | 0.0025148
27 | C3 | 5 | SMN1 | 0.47883 | -0.00060965
27 | C3 | 6 | SMN1 | 0.47844 | -0.0015528
27 | C3 | 7 | SMN1 | 0.47979 | -0.00075727
27 | C3 | 8 | SMN1 | 0.48055 | -0.00055219
27 | C3 | 9 | SMN1 | 0.48232 | 0.00066804
27 | C3 | 10 | SMN1 | 0.48282 | 0.00061074
27 | C3 | 11 | SMN1 | 0.48283 | 6.8768e-05
27 | C3 | 12 | SMN1 | 0.4811 | -0.0022115
27 | C3 | 13 | SMN1 | 0.48243 | -0.0014385
27 | C3 | 14 | SMN1 | 0.48427 | -0.00015511
27 | C3 | 15 | SMN1 | 0.48229 | -0.0026864
27 | C3 | 16 | SMN1 | 0.4826 | -0.0029301
27 | C3 | 17 | SMN1 | 0.48587 | -0.00021856
27 | C3 | 18 | SMN1 | 0.48939 | 0.0027503
27 | C3 | 19 | SMN1 | 0.49147 | 0.0042779
27 | C3 | 20 | SMN1 | 0.49899 | 0.011236
27 | C3 | 21 | SMN1 | 0.5112 | 0.022898
27 | C3 | 22 | SMN1 | 0.53689 | 0.048034
27 | C3 | 23 | SMN1 | 0.58486 | 0.095446
27 | C3 | 24 | SMN1 | 0.66146 | 0.17149
27 | C3 | 25 | SMN1 | 0.75412 | 0.2636
27 | C3 | 26 | SMN1 | 0.8621 | 0.37103
27 | C3 | 27 | SMN1 | 0.97033 | 0.4787
27 | C3 | 28 | SMN1 | 1.0753 | 0.58309
27 | C3 | 29 | SMN1 | 1.177 | 0.68431
27 | C3 | 30 | SMN1 | 1.2847 | 0.79137
27 | C3 | 31 | SMN1 | 1.3935 | 0.89964
27 | C3 | 32 | SMN1 | 1.4935 | 0.99909
27 | C3 | 33 | SMN1 | 1.5829 | 1.088
27 | C3 | 34 | SMN1 | 1.677 | 1.1814
27 | C3 | 35 | SMN1 | 1.7538 | 1.2577
28 | C4 | 1
28 | C4 | 2
28 | C4 | 3
28 | C4 | 4
28 | C4 | 5
28 | C4 | 6
28 | C4 | 7
28 | C4 | 8
28 | C4 | 9
28 | C4 | 10
28 | C4 | 11
28 | C4 | 12
28 | C4 | 13
28 | C4 | 14
28 | C4 | 15
28 | C4 | 16
28 | C4 | 17
28 | C4 | 18
28 | C4 | 19
28 | C4 | 20
28 | C4 | 21
28 | C4 | 22
28 | C4 | 23
28 | C4 | 24
28 | C4 | 25
28 | C4 | 26
28 | C4 | 27
28 | C4 | 28
28 | C4 | 29
28 | C4 | 30
28 | C4 | 31
28 | C4 | 32
28 | C4 | 33
28 | C4 | 34
28 | C4 | 35
29 | C5 | 1
29 | C5 | 2
29 | C5 | 3
29 | C5 | 4
29 | C5 | 5
29 | C5 | 6
29 | C5 | 7
29 | C5 | 8
29 | C5 | 9
29 | C5 | 10
29 | C5 | 11
29 | C5 | 12
29 | C5 | 13
29 | C5 | 14
29 | C5 | 15
29 | C5 | 16
29 | C5 | 17
29 | C5 | 18
29 | C5 | 19
29 | C5 | 20
29 | C5 | 21
29 | C5 | 22
29 | C5 | 23
29 | C5 | 24
29 | C5 | 25
29 | C5 | 26
29 | C5 | 27
29 | C5 | 28
29 | C5 | 29
29 | C5 | 30
29 | C5 | 31
29 | C5 | 32
29 | C5 | 33
29 | C5 | 34
29 | C5 | 35
30 | C6 | 1
30 | C6 | 2
30 | C6 | 3
30 | C6 | 4
30 | C6 | 5
30 | C6 | 6
30 | C6 | 7
30 | C6 | 8
30 | C6 | 9
30 | C6 | 10
30 | C6 | 11
30 | C6 | 12
30 | C6 | 13
30 | C6 | 14
30 | C6 | 15
30 | C6 | 16
30 | C6 | 17
30 | C6 | 18
30 | C6 | 19
30 | C6 | 20
30 | C6 | 21
30 | C6 | 22
30 | C6 | 23
30 | C6 | 24
30 | C6 | 25
30 | C6 | 26
30 | C6 | 27
30 | C6 | 28
30 | C6 | 29
30 | C6 | 30
30 | C6 | 31
30 | C6 | 32
30 | C6 | 33
30 | C6 | 34
30 | C6 | 35
31 | C7 | 1
31 | C7 | 2
31 | C7 | 3
31 | C7 | 4
31 | C7 | 5
31 | C7 | 6
31 | C7 | 7
31 | C7 | 8
31 | C7 | 9
31 | C7 | 10
31 | C7 | 11
31 | C7 | 12
31 | C7 | 13
31 | C7 | 14
31 | C7 | 15
31 | C7 | 16
31 | C7 | 17
31 | C7 | 18
31 | C7 | 19
31 | C7 | 20
31 | C7 | 21
31 | C7 | 22
31 | C7 | 23
31 | C7 | 24
31 | C7 | 25
31 | C7 | 26
31 | C7 | 27
31 | C7 | 28
31 | C7 | 29
31 | C7 | 30
31 | C7 | 31
31 | C7 | 32
31 | C7 | 33
31 | C7 | 34
31 | C7 | 35
32 | C8 | 1
32 | C8 | 2
32 | C8 | 3
32 | C8 | 4
32 | C8 | 5
32 | C8 | 6
32 | C8 | 7
32 | C8 | 8
32 | C8 | 9
32 | C8 | 10
32 | C8 | 11
32 | C8 | 12
32 | C8 | 13
32 | C8 | 14
32 | C8 | 15
32 | C8 | 16
32 | C8 | 17
32 | C8 | 18
32 | C8 | 19
32 | C8 | 20
32 | C8 | 21
32 | C8 | 22
32 | C8 | 23
32 | C8 | 24
32 | C8 | 25
32 | C8 | 26
32 | C8 | 27
32 | C8 | 28
32 | C8 | 29
32 | C8 | 30
32 | C8 | 31
32 | C8 | 32
32 | C8 | 33
32 | C8 | 34
32 | C8 | 35
33 | C9 | 1
33 | C9 | 2
33 | C9 | 3
33 | C9 | 4
33 | C9 | 5
33 | C9 | 6
33 | C9 | 7
33 | C9 | 8
33 | C9 | 9
33 | C9 | 10
33 | C9 | 11
33 | C9 | 12
33 | C9 | 13
33 | C9 | 14
33 | C9 | 15
33 | C9 | 16
33 | C9 | 17
33 | C9 | 18
33 | C9 | 19
33 | C9 | 20
33 | C9 | 21
33 | C9 | 22
33 | C9 | 23
33 | C9 | 24
33 | C9 | 25
33 | C9 | 26
33 | C9 | 27
33 | C9 | 28
33 | C9 | 29
33 | C9 | 30
33 | C9 | 31
33 | C9 | 32
33 | C9 | 33
33 | C9 | 34
33 | C9 | 35
34 | C10 | 1
34 | C10 | 2
34 | C10 | 3
34 | C10 | 4
34 | C10 | 5
34 | C10 | 6
34 | C10 | 7
34 | C10 | 8
34 | C10 | 9
34 | C10 | 10
34 | C10 | 11
34 | C10 | 12
34 | C10 | 13
34 | C10 | 14
34 | C10 | 15
34 | C10 | 16
34 | C10 | 17
34 | C10 | 18
34 | C10 | 19
34 | C10 | 20
34 | C10 | 21
34 | C10 | 22
34 | C10 | 23
34 | C10 | 24
34 | C10 | 25
34 | C10 | 26
34 | C10 | 27
34 | C10 | 28
34 | C10 | 29
34 | C10 | 30
34 | C10 | 31
34 | C10 | 32
34 | C10 | 33
34 | C10 | 34
34 | C10 | 35
35 | C11 | 1
35 | C11 | 2
35 | C11 | 3
35 | C11 | 4
35 | C11 | 5
35 | C11 | 6
35 | C11 | 7
35 | C11 | 8
35 | C11 | 9
35 | C11 | 10
35 | C11 | 11
35 | C11 | 12
35 | C11 | 13
35 | C11 | 14
35 | C11 | 15
35 | C11 | 16
35 | C11 | 17
35 | C11 | 18
35 | C11 | 19
35 | C11 | 20
35 | C11 | 21
35 | C11 | 22
35 | C11 | 23
35 | C11 | 24
35 | C11 | 25
35 | C11 | 26
35 | C11 | 27
35 | C11 | 28
35 | C11 | 29
35 | C11 | 30
35 | C11 | 31
35 | C11 | 32
35 | C11 | 33
35 | C11 | 34
35 | C11 | 35
36 | C12 | 1
36 | C12 | 2
36 | C12 | 3
36 | C12 | 4
36 | C12 | 5
36 | C12 | 6
36 | C12 | 7
36 | C12 | 8
36 | C12 | 9
36 | C12 | 10
36 | C12 | 11
36 | C12 | 12
36 | C12 | 13
36 | C12 | 14
36 | C12 | 15
36 | C12 | 16
36 | C12 | 17
36 | C12 | 18
36 | C12 | 19
36 | C12 | 20
36 | C12 | 21
36 | C12 | 22
36 | C12 | 23
36 | C12 | 24
36 | C12 | 25
36 | C12 | 26
36 | C12 | 27
36 | C12 | 28
36 | C12 | 29
36 | C12 | 30
36 | C12 | 31
36 | C12 | 32
36 | C12 | 33
36 | C12 | 34
36 | C12 | 35
37 | D1 | 1
37 | D1 | 2
37 | D1 | 3
37 | D1 | 4
37 | D1 | 5
37 | D1 | 6
37 | D1 | 7
37 | D1 | 8
37 | D1 | 9
37 | D1 | 10
37 | D1 | 11
37 | D1 | 12
37 | D1 | 13
37 | D1 | 14
37 | D1 | 15
37 | D1 | 16
37 | D1 | 17
37 | D1 | 18
37 | D1 | 19
37 | D1 | 20
37 | D1 | 21
37 | D1 | 22
37 | D1 | 23
37 | D1 | 24
37 | D1 | 25
37 | D1 | 26
37 | D1 | 27
37 | D1 | 28
37 | D1 | 29
37 | D1 | 30
37 | D1 | 31
37 | D1 | 32
37 | D1 | 33
37 | D1 | 34
37 | D1 | 35
38 | D2 | 1
38 | D2 | 2
38 | D2 | 3
38 | D2 | 4
38 | D2 | 5
38 | D2 | 6
38 | D2 | 7
38 | D2 | 8
38 | D2 | 9
38 | D2 | 10
38 | D2 | 11
38 | D2 | 12
38 | D2 | 13
38 | D2 | 14
38 | D2 | 15
38 | D2 | 16
38 | D2 | 17
38 | D2 | 18
38 | D2 | 19
38 | D2 | 20
38 | D2 | 21
38 | D2 | 22
38 | D2 | 23
38 | D2 | 24
38 | D2 | 25
38 | D2 | 26
38 | D2 | 27
38 | D2 | 28
38 | D2 | 29
38 | D2 | 30
38 | D2 | 31
38 | D2 | 32
38 | D2 | 33
38 | D2 | 34
38 | D2 | 35
39 | D3 | 1 | RNP | 0.091357 | 0.00097513
39 | D3 | 2 | RNP | 0.091401 | 0.00093635
39 | D3 | 3 | RNP | 0.092653 | 0.002106
39 | D3 | 4 | RNP | 0.09244 | 0.0018097
39 | D3 | 5 | RNP | 0.091355 | 0.00064213
39 | D3 | 6 | RNP | 0.090262 | -0.00053277
39 | D3 | 7 | RNP | 0.090407 | -0.00047049
39 | D3 | 8 | RNP | 0.092008 | 0.0010475
39 | D3 | 9 | RNP | 0.091873 | 0.00082943
39 | D3 | 10 | RNP | 0.089856 | -0.0012695
39 | D3 | 11 | RNP | 0.088098 | -0.0031108
39 | D3 | 12 | RNP | 0.088537 | -0.0027545
39 | D3 | 13 | RNP | 0.09008 | -0.0012942
39 | D3 | 14 | RNP | 0.090782 | -0.00067443
39 | D3 | 15 | RNP | 0.091556 | 1.6989e-05
39 | D3 | 16 | RNP | 0.091464 | -0.00015851
39 | D3 | 17 | RNP | 0.091277 | -0.00042757
39 | D3 | 18 | RNP | 0.091411 | -0.00037699
39 | D3 | 19 | RNP | 0.091563 | -0.00030723
39 | D3 | 20 | RNP | 0.093097 | 0.0011438
39 | D3 | 21 | RNP | 0.095817 | 0.0037814
39 | D3 | 22 | RNP | 0.099242 | 0.0071237
39 | D3 | 23 | RNP | 0.10777 | 0.015572
39 | D3 | 24 | RNP | 0.12516 | 0.032881
39 | D3 | 25 | RNP | 0.15454 | 0.062178
39 | D3 | 26 | RNP | 0.19684 | 0.1044
39 | D3 | 27 | RNP | 0.24657 | 0.15404
39 | D3 | 28 | RNP | 0.29977 | 0.20716
39 | D3 | 29 | RNP | 0.35073 | 0.25803
39 | D3 | 30 | RNP | 0.39855 | 0.30577
39 | D3 | 31 | RNP | 0.43802 | 0.34516
39 | D3 | 32 | RNP | 0.47198 | 0.37904
39 | D3 | 33 | RNP | 0.50218 | 0.40915
39 | D3 | 34 | RNP | 0.52634 | 0.43323
39 | D3 | 35 | RNP | 0.54301 | 0.44981
39 | D3 | 1 | SMN2 | 0.32415 | 0.0049725
39 | D3 | 2 | SMN2 | 0.32389 | 0.0045737
39 | D3 | 3 | SMN2 | 0.3231 | 0.0036531
39 | D3 | 4 | SMN2 | 0.32174 | 0.0021617
39 | D3 | 5 | SMN2 | 0.32004 | 0.00032546
39 | D3 | 6 | SMN2 | 0.31958 | -0.0002734
39 | D3 | 7 | SMN2 | 0.31978 | -0.00020254
39 | D3 | 8 | SMN2 | 0.31933 | -0.00079303
39 | D3 | 9 | SMN2 | 0.31807 | -0.0021846
39 | D3 | 10 | SMN2 | 0.31787 | -0.0025186
39 | D3 | 11 | SMN2 | 0.31899 | -0.0015355
39 | D3 | 12 | SMN2 | 0.32039 | -0.00026657
39 | D3 | 13 | SMN2 | 0.32088 | 8.9971e-05
39 | D3 | 14 | SMN2 | 0.32143 | 0.00050522
39 | D3 | 15 | SMN2 | 0.32036 | -0.00069606
39 | D3 | 16 | SMN2 | 0.31875 | -0.0024492
39 | D3 | 17 | SMN2 | 0.31908 | -0.0022494
39 | D3 | 18 | SMN2 | 0.3213 | -0.00016539
39 | D3 | 19 | SMN2 | 0.32231 | 0.00071332
39 | D3 | 20 | SMN2 | 0.32317 | 0.0014409
39 | D3 | 21 | SMN2 | 0.32631 | 0.0044446
39 | D3 | 22 | SMN2 | 0.33071 | 0.0087057
39 | D3 | 23 | SMN2 | 0.34418 | 0.02205
39 | D3 | 24 | SMN2 | 0.3682 | 0.045935
39 | D3 | 25 | SMN2 | 0.42075 | 0.098351
39 | D3 | 26 | SMN2 | 0.49628 | 0.17374
39 | D3 | 27 | SMN2 | 0.59526 | 0.27259
39 | D3 | 28 | SMN2 | 0.69634 | 0.37353
39 | D3 | 29 | SMN2 | 0.7941 | 0.47116
39 | D3 | 30 | SMN2 | 0.89043 | 0.56736
39 | D3 | 31 | SMN2 | 0.97903 | 0.65582
39 | D3 | 32 | SMN2 | 1.0608 | 0.73741
39 | D3 | 33 | SMN2 | 1.1388 | 0.8153
39 | D3 | 34 | SMN2 | 1.2135 | 0.88989
39 | D3 | 35 | SMN2 | 1.2749 | 0.95111
39 | D3 | 1 | SMN1 | 0.49715 | 0.0025596
39 | D3 | 2 | SMN1 | 0.49748 | 0.002601
39 | D3 | 3 | SMN1 | 0.49684 | 0.0016734
39 | D3 | 4 | SMN1 | 0.49646 | 0.0010055
39 | D3 | 5 | SMN1 | 0.49577 | 2.7559e-05
39 | D3 | 6 | SMN1 | 0.49468 | -0.001353
39 | D3 | 7 | SMN1 | 0.49612 | -0.00020158
39 | D3 | 8 | SMN1 | 0.49829 | 0.0016803
39 | D3 | 9 | SMN1 | 0.4954 | -0.0014922
39 | D3 | 10 | SMN1 | 0.4948 | -0.0023841
39 | D3 | 11 | SMN1 | 0.49626 | -0.0012103
39 | D3 | 12 | SMN1 | 0.49881 | 0.0010476
39 | D3 | 13 | SMN1 | 0.4996 | 0.0015557
39 | D3 | 14 | SMN1 | 0.49755 | -0.00078559
39 | D3 | 15 | SMN1 | 0.49763 | -0.00099093
39 | D3 | 16 | SMN1 | 0.49823 | -0.0006835
39 | D3 | 17 | SMN1 | 0.49858 | -0.0006234
39 | D3 | 18 | SMN1 | 0.49935 | -0.00013629
39 | D3 | 19 | SMN1 | 0.50265 | 0.0028708
39 | D3 | 20 | SMN1 | 0.50968 | 0.0096102
39 | D3 | 21 | SMN1 | 0.52272 | 0.022368
39 | D3 | 22 | SMN1 | 0.54879 | 0.04815
39 | D3 | 23 | SMN1 | 0.59197 | 0.091044
39 | D3 | 24 | SMN1 | 0.66655 | 0.16533
39 | D3 | 25 | SMN1 | 0.76893 | 0.26742
39 | D3 | 26 | SMN1 | 0.88728 | 0.38549
39 | D3 | 27 | SMN1 | 1.0127 | 0.51065
39 | D3 | 28 | SMN1 | 1.1329 | 0.63051
39 | D3 | 29 | SMN1 | 1.2467 | 0.74404
39 | D3 | 30 | SMN1 | 1.364 | 0.86109
39 | D3 | 31 | SMN1 | 1.4765 | 0.97326
39 | D3 | 32 | SMN1 | 1.5854 | 1.0819
39 | D3 | 33 | SMN1 | 1.7013 | 1.1975
39 | D3 | 34 | SMN1 | 1.8193 | 1.3152
39 | D3 | 35 | SMN1 | 1.9132 | 1.4088
40 | D4 | 1
40 | D4 | 2
40 | D4 | 3
40 | D4 | 4
40 | D4 | 5
40 | D4 | 6
40 | D4 | 7
40 | D4 | 8
40 | D4 | 9
40 | D4 | 10
40 | D4 | 11
40 | D4 | 12
40 | D4 | 13
40 | D4 | 14
40 | D4 | 15
40 | D4 | 16
40 | D4 | 17
40 | D4 | 18
40 | D4 | 19
40 | D4 | 20
40 | D4 | 21
40 | D4 | 22
40 | D4 | 23
40 | D4 | 24
40 | D4 | 25
40 | D4 | 26
40 | D4 | 27
40 | D4 | 28
40 | D4 | 29
40 | D4 | 30
40 | D4 | 31
40 | D4 | 32
40 | D4 | 33
40 | D4 | 34
40 | D4 | 35
41 | D5 | 1 | RNP | 0.10161 | 0.0012117
41 | D5 | 2 | RNP | 0.10205 | 0.001694
41 | D5 | 3 | RNP | 0.10123 | 0.00091859
41 | D5 | 4 | RNP | 0.10053 | 0.00025873
41 | D5 | 5 | RNP | 0.099762 | -0.00047086
41 | D5 | 6 | RNP | 0.099577 | -0.0006143
41 | D5 | 7 | RNP | 0.10113 | 0.00098381
41 | D5 | 8 | RNP | 0.10057 | 0.00045889
41 | D5 | 9 | RNP | 0.10001 | -6.2082e-05
41 | D5 | 10 | RNP | 0.10043 | 0.0004043
41 | D5 | 11 | RNP | 0.10033 | 0.00035007
41 | D5 | 12 | RNP | 0.10097 | 0.0010231
41 | D5 | 13 | RNP | 0.099806 | -9.4897e-05
41 | D5 | 14 | RNP | 0.10014 | 0.00028417
41 | D5 | 15 | RNP | 0.10258 | 0.0027644
41 | D5 | 16 | RNP | 0.10022 | 0.00044698
41 | D5 | 17 | RNP | 0.099034 | -0.00070075
41 | D5 | 18 | RNP | 0.098378 | -0.0013161
41 | D5 | 19 | RNP | 0.09847 | -0.0011824
41 | D5 | 20 | RNP | 0.099636 | 2.5453e-05
41 | D5 | 21 | RNP | 0.09853 | -0.0010389
41 | D5 | 22 | RNP | 0.098228 | -0.0013001
41 | D5 | 23 | RNP | 0.099214 | -0.00027248
41 | D5 | 24 | RNP | 0.098681 | -0.00076409
41 | D5 | 25 | RNP | 0.098674 | -0.00072923
41 | D5 | 26 | RNP | 0.097884 | -0.0014781
41 | D5 | 27 | RNP | 0.097487 | -0.0018335
41 | D5 | 28 | RNP | 0.097609 | -0.0016693
41 | D5 | 29 | RNP | 0.09698 | -0.002257
41 | D5 | 30 | RNP | 0.0972 | -0.0019956
41 | D5 | 31 | RNP | 0.098498 | -0.00065594
41 | D5 | 32 | RNP | 0.099767 | 0.00065486
41 | D5 | 33 | RNP | 0.10348 | 0.0044108
41 | D5 | 34 | RNP | 0.10448 | 0.0054515
41 | D5 | 35 | RNP | 0.10708 | 0.0080906
41 | D5 | 1 | SMN2 | 0.31269 | 0.0017752
41 | D5 | 2 | SMN2 | 0.31378 | 0.0027912
41 | D5 | 3 | SMN2 | 0.31224 | 0.0011825
41 | D5 | 4 | SMN2 | 0.31108 | -5.3109e-05
41 | D5 | 5 | SMN2 | 0.31032 | -0.00088848
41 | D5 | 6 | SMN2 | 0.3116 | 0.00031188
41 | D5 | 7 | SMN2 | 0.3125 | 0.0011425
41 | D5 | 8 | SMN2 | 0.31349 | 0.0020577
41 | D5 | 9 | SMN2 | 0.31285 | 0.0013346
41 | D5 | 10 | SMN2 | 0.3116 | 1.8305e-05
41 | D5 | 11 | SMN2 | 0.30983 | -0.0018287
41 | D5 | 12 | SMN2 | 0.30878 | -0.0029536
41 | D5 | 13 | SMN2 | 0.30934 | -0.0024695
41 | D5 | 14 | SMN2 | 0.31168 | -0.00020619
41 | D5 | 15 | SMN2 | 0.31228 | 0.00032195
41 | D5 | 16 | SMN2 | 0.31267 | 0.0006275
41 | D5 | 17 | SMN2 | 0.31289 | 0.00077343
41 | D5 | 18 | SMN2 | 0.31239 | 0.00020506
41 | D5 | 19 | SMN2 | 0.31109 | -0.0011762
41 | D5 | 20 | SMN2 | 0.31089 | -0.0014527
41 | D5 | 21 | SMN2 | 0.31184 | -0.00057458
41 | D5 | 22 | SMN2 | 0.31284 | 0.00034638
41 | D5 | 23 | SMN2 | 0.31289 | 0.00032128
41 | D5 | 24 | SMN2 | 0.31123 | -0.0014121
41 | D5 | 25 | SMN2 | 0.31287 | 0.00015978
41 | D5 | 26 | SMN2 | 0.31431 | 0.0015203
41 | D5 | 27 | SMN2 | 0.31483 | 0.0019666
41 | D5 | 28 | SMN2 | 0.31448 | 0.001545
41 | D5 | 29 | SMN2 | 0.31416 | 0.0011487
41 | D5 | 30 | SMN2 | 0.31412 | 0.0010332
41 | D5 | 31 | SMN2 | 0.31359 | 0.00042485
41 | D5 | 32 | SMN2 | 0.3124 | -0.00084233
41 | D5 | 33 | SMN2 | 0.31209 | -0.0012208
41 | D5 | 34 | SMN2 | 0.31203 | -0.0013631
41 | D5 | 35 | SMN2 | 0.31199 | -0.0014794
41 | D5 | 1 | SMN1 | 0.44862 | 0.0044725
41 | D5 | 2 | SMN1 | 0.44897 | 0.0044253
41 | D5 | 3 | SMN1 | 0.44727 | 0.0023362
41 | D5 | 4 | SMN1 | 0.44848 | 0.0031471
41 | D5 | 5 | SMN1 | 0.44655 | 0.00082139
41 | D5 | 6 | SMN1 | 0.44543 | -0.00068876
41 | D5 | 7 | SMN1 | 0.44667 | 0.00015516
41 | D5 | 8 | SMN1 | 0.44735 | 0.00043353
41 | D5 | 9 | SMN1 | 0.44691 | -0.00039594
41 | D5 | 10 | SMN1 | 0.44732 | -0.00038537
41 | D5 | 11 | SMN1 | 0.44784 | -0.00025583
41 | D5 | 12 | SMN1 | 0.45074 | 0.0022433
41 | D5 | 13 | SMN1 | 0.44897 | 8.3795e-05
41 | D5 | 14 | SMN1 | 0.44951 | 0.00022367
41 | D5 | 15 | SMN1 | 0.45093 | 0.0012575
41 | D5 | 16 | SMN1 | 0.44884 | -0.0012264
41 | D5 | 17 | SMN1 | 0.44844 | -0.002027
41 | D5 | 18 | SMN1 | 0.44934 | -0.0015224
41 | D5 | 19 | SMN1 | 0.44911 | -0.0021449
41 | D5 | 20 | SMN1 | 0.44961 | -0.0020375
41 | D5 | 21 | SMN1 | 0.45207 | 2.3842e-05
41 | D5 | 22 | SMN1 | 0.45178 | -0.00065903
41 | D5 | 23 | SMN1 | 0.45085 | -0.0019881
41 | D5 | 24 | SMN1 | 0.44985 | -0.0033812
41 | D5 | 25 | SMN1 | 0.45472 | 0.0010938
41 | D5 | 26 | SMN1 | 0.45545 | 0.0014251
41 | D5 | 27 | SMN1 | 0.44914 | -0.005273
41 | D5 | 28 | SMN1 | 0.44891 | -0.0058962
41 | D5 | 29 | SMN1 | 0.45329 | -0.0019179
41 | D5 | 30 | SMN1 | 0.45759 | 0.0019921
41 | D5 | 31 | SMN1 | 0.45845 | 0.0024552
41 | D5 | 32 | SMN1 | 0.45912 | 0.0027356
41 | D5 | 33 | SMN1 | 0.4602 | 0.0034203
41 | D5 | 34 | SMN1 | 0.46313 | 0.0059522
41 | D5 | 35 | SMN1 | 0.46847 | 0.010896
42 | D6 | 1
42 | D6 | 2
42 | D6 | 3
42 | D6 | 4
42 | D6 | 5
42 | D6 | 6
42 | D6 | 7
42 | D6 | 8
42 | D6 | 9
42 | D6 | 10
42 | D6 | 11
42 | D6 | 12
42 | D6 | 13
42 | D6 | 14
42 | D6 | 15
42 | D6 | 16
42 | D6 | 17
42 | D6 | 18
42 | D6 | 19
42 | D6 | 20
42 | D6 | 21
42 | D6 | 22
42 | D6 | 23
42 | D6 | 24
42 | D6 | 25
42 | D6 | 26
42 | D6 | 27
42 | D6 | 28
42 | D6 | 29
42 | D6 | 30
42 | D6 | 31
42 | D6 | 32
42 | D6 | 33
42 | D6 | 34
42 | D6 | 35
43 | D7 | 1
43 | D7 | 2
43 | D7 | 3
43 | D7 | 4
43 | D7 | 5
43 | D7 | 6
43 | D7 | 7
43 | D7 | 8
43 | D7 | 9
43 | D7 | 10
43 | D7 | 11
43 | D7 | 12
43 | D7 | 13
43 | D7 | 14
43 | D7 | 15
43 | D7 | 16
43 | D7 | 17
43 | D7 | 18
43 | D7 | 19
43 | D7 | 20
43 | D7 | 21
43 | D7 | 22
43 | D7 | 23
43 | D7 | 24
43 | D7 | 25
43 | D7 | 26
43 | D7 | 27
43 | D7 | 28
43 | D7 | 29
43 | D7 | 30
43 | D7 | 31
43 | D7 | 32
43 | D7 | 33
43 | D7 | 34
43 | D7 | 35
44 | D8 | 1
44 | D8 | 2
44 | D8 | 3
44 | D8 | 4
44 | D8 | 5
44 | D8 | 6
44 | D8 | 7
44 | D8 | 8
44 | D8 | 9
44 | D8 | 10
44 | D8 | 11
44 | D8 | 12
44 | D8 | 13
44 | D8 | 14
44 | D8 | 15
44 | D8 | 16
44 | D8 | 17
44 | D8 | 18
44 | D8 | 19
44 | D8 | 20
44 | D8 | 21
44 | D8 | 22
44 | D8 | 23
44 | D8 | 24
44 | D8 | 25
44 | D8 | 26
44 | D8 | 27
44 | D8 | 28
44 | D8 | 29
44 | D8 | 30
44 | D8 | 31
44 | D8 | 32
44 | D8 | 33
44 | D8 | 34
44 | D8 | 35
45 | D9 | 1
45 | D9 | 2
45 | D9 | 3
45 | D9 | 4
45 | D9 | 5
45 | D9 | 6
45 | D9 | 7
45 | D9 | 8
45 | D9 | 9
45 | D9 | 10
45 | D9 | 11
45 | D9 | 12
45 | D9 | 13
45 | D9 | 14
45 | D9 | 15
45 | D9 | 16
45 | D9 | 17
45 | D9 | 18
45 | D9 | 19
45 | D9 | 20
45 | D9 | 21
45 | D9 | 22
45 | D9 | 23
45 | D9 | 24
45 | D9 | 25
45 | D9 | 26
45 | D9 | 27
45 | D9 | 28
45 | D9 | 29
45 | D9 | 30
45 | D9 | 31
45 | D9 | 32
45 | D9 | 33
45 | D9 | 34
45 | D9 | 35
46 | D10 | 1
46 | D10 | 2
46 | D10 | 3
46 | D10 | 4
46 | D10 | 5
46 | D10 | 6
46 | D10 | 7
46 | D10 | 8
46 | D10 | 9
46 | D10 | 10
46 | D10 | 11
46 | D10 | 12
46 | D10 | 13
46 | D10 | 14
46 | D10 | 15
46 | D10 | 16
46 | D10 | 17
46 | D10 | 18
46 | D10 | 19
46 | D10 | 20
46 | D10 | 21
46 | D10 | 22
46 | D10 | 23
46 | D10 | 24
46 | D10 | 25
46 | D10 | 26
46 | D10 | 27
46 | D10 | 28
46 | D10 | 29
46 | D10 | 30
46 | D10 | 31
46 | D10 | 32
46 | D10 | 33
46 | D10 | 34
46 | D10 | 35
47 | D11 | 1
47 | D11 | 2
47 | D11 | 3
47 | D11 | 4
47 | D11 | 5
47 | D11 | 6
47 | D11 | 7
47 | D11 | 8
47 | D11 | 9
47 | D11 | 10
47 | D11 | 11
47 | D11 | 12
47 | D11 | 13
47 | D11 | 14
47 | D11 | 15
47 | D11 | 16
47 | D11 | 17
47 | D11 | 18
47 | D11 | 19
47 | D11 | 20
47 | D11 | 21
47 | D11 | 22
47 | D11 | 23
47 | D11 | 24
47 | D11 | 25
47 | D11 | 26
47 | D11 | 27
47 | D11 | 28
47 | D11 | 29
47 | D11 | 30
47 | D11 | 31
47 | D11 | 32
47 | D11 | 33
47 | D11 | 34
47 | D11 | 35
48 | D12 | 1
48 | D12 | 2
48 | D12 | 3
48 | D12 | 4
48 | D12 | 5
48 | D12 | 6
48 | D12 | 7
48 | D12 | 8
48 | D12 | 9
48 | D12 | 10
48 | D12 | 11
48 | D12 | 12
48 | D12 | 13
48 | D12 | 14
48 | D12 | 15
48 | D12 | 16
48 | D12 | 17
48 | D12 | 18
48 | D12 | 19
48 | D12 | 20
48 | D12 | 21
48 | D12 | 22
48 | D12 | 23
48 | D12 | 24
48 | D12 | 25
48 | D12 | 26
48 | D12 | 27
48 | D12 | 28
48 | D12 | 29
48 | D12 | 30
48 | D12 | 31
48 | D12 | 32
48 | D12 | 33
48 | D12 | 34
48 | D12 | 35
49 | E1 | 1
49 | E1 | 2
49 | E1 | 3
49 | E1 | 4
49 | E1 | 5
49 | E1 | 6
49 | E1 | 7
49 | E1 | 8
49 | E1 | 9
49 | E1 | 10
49 | E1 | 11
49 | E1 | 12
49 | E1 | 13
49 | E1 | 14
49 | E1 | 15
49 | E1 | 16
49 | E1 | 17
49 | E1 | 18
49 | E1 | 19
49 | E1 | 20
49 | E1 | 21
49 | E1 | 22
49 | E1 | 23
49 | E1 | 24
49 | E1 | 25
49 | E1 | 26
49 | E1 | 27
49 | E1 | 28
49 | E1 | 29
49 | E1 | 30
49 | E1 | 31
49 | E1 | 32
49 | E1 | 33
49 | E1 | 34
49 | E1 | 35
50 | E2 | 1
50 | E2 | 2
50 | E2 | 3
50 | E2 | 4
50 | E2 | 5
50 | E2 | 6
50 | E2 | 7
50 | E2 | 8
50 | E2 | 9
50 | E2 | 10
50 | E2 | 11
50 | E2 | 12
50 | E2 | 13
50 | E2 | 14
50 | E2 | 15
50 | E2 | 16
50 | E2 | 17
50 | E2 | 18
50 | E2 | 19
50 | E2 | 20
50 | E2 | 21
50 | E2 | 22
50 | E2 | 23
50 | E2 | 24
50 | E2 | 25
50 | E2 | 26
50 | E2 | 27
50 | E2 | 28
50 | E2 | 29
50 | E2 | 30
50 | E2 | 31
50 | E2 | 32
50 | E2 | 33
50 | E2 | 34
50 | E2 | 35
51 | E3 | 1 | RNP | 0.09857 | -0.00046481
51 | E3 | 2 | RNP | 0.098316 | -0.00075184
51 | E3 | 3 | RNP | 0.099601 | 0.00049998
51 | E3 | 4 | RNP | 0.10002 | 0.00088758
51 | E3 | 5 | RNP | 0.10154 | 0.0023757
51 | E3 | 6 | RNP | 0.10063 | 0.0014302
51 | E3 | 7 | RNP | 0.098726 | -0.00050649
51 | E3 | 8 | RNP | 0.098106 | -0.0011591
51 | E3 | 9 | RNP | 0.098238 | -0.00106
51 | E3 | 10 | RNP | 0.098806 | -0.00052536
51 | E3 | 11 | RNP | 0.099375 | 1.1074e-05
51 | E3 | 12 | RNP | 0.099471 | 7.4044e-05
51 | E3 | 13 | RNP | 0.099093 | -0.00033729
51 | E3 | 14 | RNP | 0.098132 | -0.0013309
51 | E3 | 15 | RNP | 0.097876 | -0.00162
51 | E3 | 16 | RNP | 0.097442 | -0.0020866
51 | E3 | 17 | RNP | 0.097445 | -0.0021167
51 | E3 | 18 | RNP | 0.097935 | -0.0016589
51 | E3 | 19 | RNP | 0.099501 | -0.00012581
51 | E3 | 20 | RNP | 0.10185 | 0.0021908
51 | E3 | 21 | RNP | 0.10475 | 0.0050577
51 | E3 | 22 | RNP | 0.1089 | 0.0091764
51 | E3 | 23 | RNP | 0.11593 | 0.016175
51 | E3 | 24 | RNP | 0.13356 | 0.033764
51 | E3 | 25 | RNP | 0.16627 | 0.06645
51 | E3 | 26 | RNP | 0.21343 | 0.11357
51 | E3 | 27 | RNP | 0.26074 | 0.16085
51 | E3 | 28 | RNP | 0.31578 | 0.21586
51 | E3 | 29 | RNP | 0.36884 | 0.26889
51 | E3 | 30 | RNP | 0.41292 | 0.31293
51 | E3 | 31 | RNP | 0.45278 | 0.35276
51 | E3 | 32 | RNP | 0.48366 | 0.38361
51 | E3 | 33 | RNP | 0.51293 | 0.41284
51 | E3 | 34 | RNP | 0.53627 | 0.43615
51 | E3 | 35 | RNP | 0.55125 | 0.4511
51 | E3 | 1 | SMN2 | 0.33733 | 0.005111
51 | E3 | 2 | SMN2 | 0.33642 | 0.0038283
51 | E3 | 3 | SMN2 | 0.33537 | 0.0024105
51 | E3 | 4 | SMN2 | 0.33442 | 0.001087
51 | E3 | 5 | SMN2 | 0.33381 | 9.6424e-05
51 | E3 | 6 | SMN2 | 0.33406 | -2.262e-05
51 | E3 | 7 | SMN2 | 0.33434 | -0.00011931
51 | E3 | 8 | SMN2 | 0.3346 | -0.00023421
51 | E3 | 9 | SMN2 | 0.33556 | 0.00035446
51 | E3 | 10 | SMN2 | 0.33503 | -0.00054693
51 | E3 | 11 | SMN2 | 0.33533 | -0.0006193
51 | E3 | 12 | SMN2 | 0.33582 | -0.00050714
51 | E3 | 13 | SMN2 | 0.33478 | -0.0019156
51 | E3 | 14 | SMN2 | 0.33488 | -0.0021902
51 | E3 | 15 | SMN2 | 0.33548 | -0.0019643
51 | E3 | 16 | SMN2 | 0.33616 | -0.0016524
51 | E3 | 17 | SMN2 | 0.33735 | -0.00084538
51 | E3 | 18 | SMN2 | 0.3383 | -0.00026377
51 | E3 | 19 | SMN2 | 0.34104 | 0.0021076
51 | E3 | 20 | SMN2 | 0.34414 | 0.0048252
51 | E3 | 21 | SMN2 | 0.35116 | 0.01148
51 | E3 | 22 | SMN2 | 0.36904 | 0.028978
51 | E3 | 23 | SMN2 | 0.40362 | 0.063186
51 | E3 | 24 | SMN2 | 0.46727 | 0.12646
51 | E3 | 25 | SMN2 | 0.56732 | 0.22615
51 | E3 | 26 | SMN2 | 0.68246 | 0.34091
51 | E3 | 27 | SMN2 | 0.80493 | 0.463
51 | E3 | 28 | SMN2 | 0.91647 | 0.57417
51 | E3 | 29 | SMN2 | 1.0268 | 0.68409
51 | E3 | 30 | SMN2 | 1.1255 | 0.78243
51 | E3 | 31 | SMN2 | 1.2166 | 0.87318
51 | E3 | 32 | SMN2 | 1.3044 | 0.96061
51 | E3 | 33 | SMN2 | 1.3829 | 1.0387
51 | E3 | 34 | SMN2 | 1.4564 | 1.1118
51 | E3 | 35 | SMN2 | 1.513 | 1.1681
51 | E3 | 1 | SMN1 | 0.47868 | 0.0050304
51 | E3 | 2 | SMN1 | 0.47931 | 0.0049782
51 | E3 | 3 | SMN1 | 0.47924 | 0.0042347
51 | E3 | 4 | SMN1 | 0.47841 | 0.0027239
51 | E3 | 5 | SMN1 | 0.4754 | -0.00096232
51 | E3 | 6 | SMN1 | 0.4761 | -0.00093733
51 | E3 | 7 | SMN1 | 0.47955 | 0.0018383
51 | E3 | 8 | SMN1 | 0.4799 | 0.0015051
51 | E3 | 9 | SMN1 | 0.47493 | -0.0041351
51 | E3 | 10 | SMN1 | 0.47461 | -0.0051369
51 | E3 | 11 | SMN1 | 0.47783 | -0.002593
51 | E3 | 12 | SMN1 | 0.48052 | -0.00058137
51 | E3 | 13 | SMN1 | 0.48129 | -0.00048844
51 | E3 | 14 | SMN1 | 0.48063 | -0.0018283
51 | E3 | 15 | SMN1 | 0.48197 | -0.0011593
51 | E3 | 16 | SMN1 | 0.48333 | -0.00048256
51 | E3 | 17 | SMN1 | 0.48547 | 0.00097904
51 | E3 | 18 | SMN1 | 0.49219 | 0.0070236
51 | E3 | 19 | SMN1 | 0.4955 | 0.0096612
51 | E3 | 20 | SMN1 | 0.50516 | 0.018635
51 | E3 | 21 | SMN1 | 0.53228 | 0.045081
51 | E3 | 22 | SMN1 | 0.58479 | 0.096913
51 | E3 | 23 | SMN1 | 0.669 | 0.18045
51 | E3 | 24 | SMN1 | 0.77094 | 0.28171
51 | E3 | 25 | SMN1 | 0.89146 | 0.40156
51 | E3 | 26 | SMN1 | 1.0127 | 0.52213
51 | E3 | 27 | SMN1 | 1.1321 | 0.64087
51 | E3 | 28 | SMN1 | 1.2459 | 0.75391
51 | E3 | 29 | SMN1 | 1.3554 | 0.86274
51 | E3 | 30 | SMN1 | 1.4702 | 0.97686
51 | E3 | 31 | SMN1 | 1.5817 | 1.0878
51 | E3 | 32 | SMN1 | 1.6892 | 1.1946
51 | E3 | 33 | SMN1 | 1.7918 | 1.2964
51 | E3 | 34 | SMN1 | 1.8941 | 1.3981
51 | E3 | 35 | SMN1 | 1.982 | 1.4853
52 | E4 | 1
52 | E4 | 2
52 | E4 | 3
52 | E4 | 4
52 | E4 | 5
52 | E4 | 6
52 | E4 | 7
52 | E4 | 8
52 | E4 | 9
52 | E4 | 10
52 | E4 | 11
52 | E4 | 12
52 | E4 | 13
52 | E4 | 14
52 | E4 | 15
52 | E4 | 16
52 | E4 | 17
52 | E4 | 18
52 | E4 | 19
52 | E4 | 20
52 | E4 | 21
52 | E4 | 22
52 | E4 | 23
52 | E4 | 24
52 | E4 | 25
52 | E4 | 26
52 | E4 | 27
52 | E4 | 28
52 | E4 | 29
52 | E4 | 30
52 | E4 | 31
52 | E4 | 32
52 | E4 | 33
52 | E4 | 34
52 | E4 | 35
53 | E5 | 1
53 | E5 | 2
53 | E5 | 3
53 | E5 | 4
53 | E5 | 5
53 | E5 | 6
53 | E5 | 7
53 | E5 | 8
53 | E5 | 9
53 | E5 | 10
53 | E5 | 11
53 | E5 | 12
53 | E5 | 13
53 | E5 | 14
53 | E5 | 15
53 | E5 | 16
53 | E5 | 17
53 | E5 | 18
53 | E5 | 19
53 | E5 | 20
53 | E5 | 21
53 | E5 | 22
53 | E5 | 23
53 | E5 | 24
53 | E5 | 25
53 | E5 | 26
53 | E5 | 27
53 | E5 | 28
53 | E5 | 29
53 | E5 | 30
53 | E5 | 31
53 | E5 | 32
53 | E5 | 33
53 | E5 | 34
53 | E5 | 35
54 | E6 | 1
54 | E6 | 2
54 | E6 | 3
54 | E6 | 4
54 | E6 | 5
54 | E6 | 6
54 | E6 | 7
54 | E6 | 8
54 | E6 | 9
54 | E6 | 10
54 | E6 | 11
54 | E6 | 12
54 | E6 | 13
54 | E6 | 14
54 | E6 | 15
54 | E6 | 16
54 | E6 | 17
54 | E6 | 18
54 | E6 | 19
54 | E6 | 20
54 | E6 | 21
54 | E6 | 22
54 | E6 | 23
54 | E6 | 24
54 | E6 | 25
54 | E6 | 26
54 | E6 | 27
54 | E6 | 28
54 | E6 | 29
54 | E6 | 30
54 | E6 | 31
54 | E6 | 32
54 | E6 | 33
54 | E6 | 34
54 | E6 | 35
55 | E7 | 1
55 | E7 | 2
55 | E7 | 3
55 | E7 | 4
55 | E7 | 5
55 | E7 | 6
55 | E7 | 7
55 | E7 | 8
55 | E7 | 9
55 | E7 | 10
55 | E7 | 11
55 | E7 | 12
55 | E7 | 13
55 | E7 | 14
55 | E7 | 15
55 | E7 | 16
55 | E7 | 17
55 | E7 | 18
55 | E7 | 19
55 | E7 | 20
55 | E7 | 21
55 | E7 | 22
55 | E7 | 23
55 | E7 | 24
55 | E7 | 25
55 | E7 | 26
55 | E7 | 27
55 | E7 | 28
55 | E7 | 29
55 | E7 | 30
55 | E7 | 31
55 | E7 | 32
55 | E7 | 33
55 | E7 | 34
55 | E7 | 35
56 | E8 | 1
56 | E8 | 2
56 | E8 | 3
56 | E8 | 4
56 | E8 | 5
56 | E8 | 6
56 | E8 | 7
56 | E8 | 8
56 | E8 | 9
56 | E8 | 10
56 | E8 | 11
56 | E8 | 12
56 | E8 | 13
56 | E8 | 14
56 | E8 | 15
56 | E8 | 16
56 | E8 | 17
56 | E8 | 18
56 | E8 | 19
56 | E8 | 20
56 | E8 | 21
56 | E8 | 22
56 | E8 | 23
56 | E8 | 24
56 | E8 | 25
56 | E8 | 26
56 | E8 | 27
56 | E8 | 28
56 | E8 | 29
56 | E8 | 30
56 | E8 | 31
56 | E8 | 32
56 | E8 | 33
56 | E8 | 34
56 | E8 | 35
57 | E9 | 1
57 | E9 | 2
57 | E9 | 3
57 | E9 | 4
57 | E9 | 5
57 | E9 | 6
57 | E9 | 7
57 | E9 | 8
57 | E9 | 9
57 | E9 | 10
57 | E9 | 11
57 | E9 | 12
57 | E9 | 13
57 | E9 | 14
57 | E9 | 15
57 | E9 | 16
57 | E9 | 17
57 | E9 | 18
57 | E9 | 19
57 | E9 | 20
57 | E9 | 21
57 | E9 | 22
57 | E9 | 23
57 | E9 | 24
57 | E9 | 25
57 | E9 | 26
57 | E9 | 27
57 | E9 | 28
57 | E9 | 29
57 | E9 | 30
57 | E9 | 31
57 | E9 | 32
57 | E9 | 33
57 | E9 | 34
57 | E9 | 35
58 | E10 | 1
58 | E10 | 2
58 | E10 | 3
58 | E10 | 4
58 | E10 | 5
58 | E10 | 6
58 | E10 | 7
58 | E10 | 8
58 | E10 | 9
58 | E10 | 10
58 | E10 | 11
58 | E10 | 12
58 | E10 | 13
58 | E10 | 14
58 | E10 | 15
58 | E10 | 16
58 | E10 | 17
58 | E10 | 18
58 | E10 | 19
58 | E10 | 20
58 | E10 | 21
58 | E10 | 22
58 | E10 | 23
58 | E10 | 24
58 | E10 | 25
58 | E10 | 26
58 | E10 | 27
58 | E10 | 28
58 | E10 | 29
58 | E10 | 30
58 | E10 | 31
58 | E10 | 32
58 | E10 | 33
58 | E10 | 34
58 | E10 | 35
59 | E11 | 1
59 | E11 | 2
59 | E11 | 3
59 | E11 | 4
59 | E11 | 5
59 | E11 | 6
59 | E11 | 7
59 | E11 | 8
59 | E11 | 9
59 | E11 | 10
59 | E11 | 11
59 | E11 | 12
59 | E11 | 13
59 | E11 | 14
59 | E11 | 15
59 | E11 | 16
59 | E11 | 17
59 | E11 | 18
59 | E11 | 19
59 | E11 | 20
59 | E11 | 21
59 | E11 | 22
59 | E11 | 23
59 | E11 | 24
59 | E11 | 25
59 | E11 | 26
59 | E11 | 27
59 | E11 | 28
59 | E11 | 29
59 | E11 | 30
59 | E11 | 31
59 | E11 | 32
59 | E11 | 33
59 | E11 | 34
59 | E11 | 35
60 | E12 | 1
60 | E12 | 2
60 | E12 | 3
60 | E12 | 4
60 | E12 | 5
60 | E12 | 6
60 | E12 | 7
60 | E12 | 8
60 | E12 | 9
60 | E12 | 10
60 | E12 | 11
60 | E12 | 12
60 | E12 | 13
60 | E12 | 14
60 | E12 | 15
60 | E12 | 16
60 | E12 | 17
60 | E12 | 18
60 | E12 | 19
60 | E12 | 20
60 | E12 | 21
60 | E12 | 22
60 | E12 | 23
60 | E12 | 24
60 | E12 | 25
60 | E12 | 26
60 | E12 | 27
60 | E12 | 28
60 | E12 | 29
60 | E12 | 30
60 | E12 | 31
60 | E12 | 32
60 | E12 | 33
60 | E12 | 34
60 | E12 | 35
61 | F1 | 1
61 | F1 | 2
61 | F1 | 3
61 | F1 | 4
61 | F1 | 5
61 | F1 | 6
61 | F1 | 7
61 | F1 | 8
61 | F1 | 9
61 | F1 | 10
61 | F1 | 11
61 | F1 | 12
61 | F1 | 13
61 | F1 | 14
61 | F1 | 15
61 | F1 | 16
61 | F1 | 17
61 | F1 | 18
61 | F1 | 19
61 | F1 | 20
61 | F1 | 21
61 | F1 | 22
61 | F1 | 23
61 | F1 | 24
61 | F1 | 25
61 | F1 | 26
61 | F1 | 27
61 | F1 | 28
61 | F1 | 29
61 | F1 | 30
61 | F1 | 31
61 | F1 | 32
61 | F1 | 33
61 | F1 | 34
61 | F1 | 35
62 | F2 | 1
62 | F2 | 2
62 | F2 | 3
62 | F2 | 4
62 | F2 | 5
62 | F2 | 6
62 | F2 | 7
62 | F2 | 8
62 | F2 | 9
62 | F2 | 10
62 | F2 | 11
62 | F2 | 12
62 | F2 | 13
62 | F2 | 14
62 | F2 | 15
62 | F2 | 16
62 | F2 | 17
62 | F2 | 18
62 | F2 | 19
62 | F2 | 20
62 | F2 | 21
62 | F2 | 22
62 | F2 | 23
62 | F2 | 24
62 | F2 | 25
62 | F2 | 26
62 | F2 | 27
62 | F2 | 28
62 | F2 | 29
62 | F2 | 30
62 | F2 | 31
62 | F2 | 32
62 | F2 | 33
62 | F2 | 34
62 | F2 | 35
63 | F3 | 1 | RNP | 0.10312 | 0.0016453
63 | F3 | 2 | RNP | 0.10315 | 0.0015112
63 | F3 | 3 | RNP | 0.10375 | 0.0019458
63 | F3 | 4 | RNP | 0.10212 | 0.00015528
63 | F3 | 5 | RNP | 0.10182 | -0.00030878
63 | F3 | 6 | RNP | 0.10301 | 0.00071281
63 | F3 | 7 | RNP | 0.10255 | 9.3867e-05
63 | F3 | 8 | RNP | 0.10263 | 1.1131e-05
63 | F3 | 9 | RNP | 0.1042 | 0.0014163
63 | F3 | 10 | RNP | 0.10424 | 0.0012932
63 | F3 | 11 | RNP | 0.10274 | -0.00037628
63 | F3 | 12 | RNP | 0.10237 | -0.00091201
63 | F3 | 13 | RNP | 0.10234 | -0.001101
63 | F3 | 14 | RNP | 0.10127 | -0.0023388
63 | F3 | 15 | RNP | 0.10152 | -0.0022503
63 | F3 | 16 | RNP | 0.10321 | -0.00072761
63 | F3 | 17 | RNP | 0.1037 | -0.00040326
63 | F3 | 18 | RNP | 0.10349 | -0.00077821
63 | F3 | 19 | RNP | 0.10304 | -0.0013835
63 | F3 | 20 | RNP | 0.10258 | -0.0020117
63 | F3 | 21 | RNP | 0.10362 | -0.0011389
63 | F3 | 22 | RNP | 0.10804 | 0.0031236
63 | F3 | 23 | RNP | 0.11006 | 0.0049783
63 | F3 | 24 | RNP | 0.11779 | 0.012546
63 | F3 | 25 | RNP | 0.12857 | 0.023161
63 | F3 | 26 | RNP | 0.14859 | 0.04301
63 | F3 | 27 | RNP | 0.17641 | 0.070672
63 | F3 | 28 | RNP | 0.21686 | 0.11096
63 | F3 | 29 | RNP | 0.26337 | 0.1573
63 | F3 | 30 | RNP | 0.30925 | 0.20302
63 | F3 | 31 | RNP | 0.35368 | 0.24728
63 | F3 | 32 | RNP | 0.39155 | 0.28499
63 | F3 | 33 | RNP | 0.42712 | 0.3204
63 | F3 | 34 | RNP | 0.45285 | 0.34596
63 | F3 | 35 | RNP | 0.46817 | 0.36112
63 | F3 | 1 | SMN2 | 0.33719 | 0.0030842
63 | F3 | 2 | SMN2 | 0.33564 | 0.0013335
63 | F3 | 3 | SMN2 | 0.33524 | 0.00072667
63 | F3 | 4 | SMN2 | 0.33569 | 0.00096976
63 | F3 | 5 | SMN2 | 0.33765 | 0.0027274
63 | F3 | 6 | SMN2 | 0.33664 | 0.0015145
63 | F3 | 7 | SMN2 | 0.33554 | 0.00021218
63 | F3 | 8 | SMN2 | 0.33346 | -0.0020752
63 | F3 | 9 | SMN2 | 0.3336 | -0.0021342
63 | F3 | 10 | SMN2 | 0.33574 | -0.00020172
63 | F3 | 11 | SMN2 | 0.33538 | -0.0007635
63 | F3 | 12 | SMN2 | 0.33566 | -0.00069315
63 | F3 | 13 | SMN2 | 0.33668 | 0.00012274
63 | F3 | 14 | SMN2 | 0.33639 | -0.00037482
63 | F3 | 15 | SMN2 | 0.33602 | -0.00094889
63 | F3 | 16 | SMN2 | 0.33472 | -0.0024488
63 | F3 | 17 | SMN2 | 0.33519 | -0.0021909
63 | F3 | 18 | SMN2 | 0.3363 | -0.0012835
63 | F3 | 19 | SMN2 | 0.33732 | -0.0004619
63 | F3 | 20 | SMN2 | 0.3409 | 0.0029058
63 | F3 | 21 | SMN2 | 0.34259 | 0.0043976
63 | F3 | 22 | SMN2 | 0.3462 | 0.0078016
63 | F3 | 23 | SMN2 | 0.35906 | 0.020451
63 | F3 | 24 | SMN2 | 0.38362 | 0.044807
63 | F3 | 25 | SMN2 | 0.43679 | 0.09778
63 | F3 | 26 | SMN2 | 0.52224 | 0.18303
63 | F3 | 27 | SMN2 | 0.64026 | 0.30084
63 | F3 | 28 | SMN2 | 0.7689 | 0.42927
63 | F3 | 29 | SMN2 | 0.89681 | 0.55697
63 | F3 | 30 | SMN2 | 1.0164 | 0.67632
63 | F3 | 31 | SMN2 | 1.1221 | 0.7819
63 | F3 | 32 | SMN2 | 1.2173 | 0.87686
63 | F3 | 33 | SMN2 | 1.3089 | 0.96821
63 | F3 | 34 | SMN2 | 1.3927 | 1.0518
63 | F3 | 35 | SMN2 | 1.4582 | 1.1171
63 | F3 | 1 | SMN1 | 0.4867 | 0.0015252
63 | F3 | 2 | SMN1 | 0.48753 | 0.0019052
63 | F3 | 3 | SMN1 | 0.48995 | 0.0038783
63 | F3 | 4 | SMN1 | 0.49012 | 0.003599
63 | F3 | 5 | SMN1 | 0.48733 | 0.0003625
63 | F3 | 6 | SMN1 | 0.48749 | 7.5775e-05
63 | F3 | 7 | SMN1 | 0.48959 | 0.0017195
63 | F3 | 8 | SMN1 | 0.48811 | -0.00020749
63 | F3 | 9 | SMN1 | 0.48751 | -0.0012556
63 | F3 | 10 | SMN1 | 0.48959 | 0.00037814
63 | F3 | 11 | SMN1 | 0.49048 | 0.00081105
63 | F3 | 12 | SMN1 | 0.48203 | -0.0080903
63 | F3 | 13 | SMN1 | 0.48231 | -0.0082553
63 | F3 | 14 | SMN1 | 0.48759 | -0.0034278
63 | F3 | 15 | SMN1 | 0.48952 | -0.0019409
63 | F3 | 16 | SMN1 | 0.49128 | -0.00063881
63 | F3 | 17 | SMN1 | 0.49237 | 1.0075e-05
63 | F3 | 18 | SMN1 | 0.49438 | 0.0015708
63 | F3 | 19 | SMN1 | 0.49828 | 0.0050197
63 | F3 | 20 | SMN1 | 0.5001 | 0.0063915
63 | F3 | 21 | SMN1 | 0.50771 | 0.013553
63 | F3 | 22 | SMN1 | 0.52381 | 0.029197
63 | F3 | 23 | SMN1 | 0.55591 | 0.060847
63 | F3 | 24 | SMN1 | 0.61222 | 0.11671
63 | F3 | 25 | SMN1 | 0.69508 | 0.19912
63 | F3 | 26 | SMN1 | 0.79229 | 0.29588
63 | F3 | 27 | SMN1 | 0.89806 | 0.40121
63 | F3 | 28 | SMN1 | 0.99349 | 0.49618
63 | F3 | 29 | SMN1 | 1.0941 | 0.59636
63 | F3 | 30 | SMN1 | 1.1873 | 0.68907
63 | F3 | 31 | SMN1 | 1.2757 | 0.77706
63 | F3 | 32 | SMN1 | 1.359 | 0.85992
63 | F3 | 33 | SMN1 | 1.4416 | 0.94206
63 | F3 | 34 | SMN1 | 1.5192 | 1.0192
63 | F3 | 35 | SMN1 | 1.5832 | 1.0828
64 | F4 | 1
64 | F4 | 2
64 | F4 | 3
64 | F4 | 4
64 | F4 | 5
64 | F4 | 6
64 | F4 | 7
64 | F4 | 8
64 | F4 | 9
64 | F4 | 10
64 | F4 | 11
64 | F4 | 12
64 | F4 | 13
64 | F4 | 14
64 | F4 | 15
64 | F4 | 16
64 | F4 | 17
64 | F4 | 18
64 | F4 | 19
64 | F4 | 20
64 | F4 | 21
64 | F4 | 22
64 | F4 | 23
64 | F4 | 24
64 | F4 | 25
64 | F4 | 26
64 | F4 | 27
64 | F4 | 28
64 | F4 | 29
64 | F4 | 30
64 | F4 | 31
64 | F4 | 32
64 | F4 | 33
64 | F4 | 34
64 | F4 | 35
65 | F5 | 1
65 | F5 | 2
65 | F5 | 3
65 | F5 | 4
65 | F5 | 5
65 | F5 | 6
65 | F5 | 7
65 | F5 | 8
65 | F5 | 9
65 | F5 | 10
65 | F5 | 11
65 | F5 | 12
65 | F5 | 13
65 | F5 | 14
65 | F5 | 15
65 | F5 | 16
65 | F5 | 17
65 | F5 | 18
65 | F5 | 19
65 | F5 | 20
65 | F5 | 21
65 | F5 | 22
65 | F5 | 23
65 | F5 | 24
65 | F5 | 25
65 | F5 | 26
65 | F5 | 27
65 | F5 | 28
65 | F5 | 29
65 | F5 | 30
65 | F5 | 31
65 | F5 | 32
65 | F5 | 33
65 | F5 | 34
65 | F5 | 35
66 | F6 | 1
66 | F6 | 2
66 | F6 | 3
66 | F6 | 4
66 | F6 | 5
66 | F6 | 6
66 | F6 | 7
66 | F6 | 8
66 | F6 | 9
66 | F6 | 10
66 | F6 | 11
66 | F6 | 12
66 | F6 | 13
66 | F6 | 14
66 | F6 | 15
66 | F6 | 16
66 | F6 | 17
66 | F6 | 18
66 | F6 | 19
66 | F6 | 20
66 | F6 | 21
66 | F6 | 22
66 | F6 | 23
66 | F6 | 24
66 | F6 | 25
66 | F6 | 26
66 | F6 | 27
66 | F6 | 28
66 | F6 | 29
66 | F6 | 30
66 | F6 | 31
66 | F6 | 32
66 | F6 | 33
66 | F6 | 34
66 | F6 | 35
67 | F7 | 1
67 | F7 | 2
67 | F7 | 3
67 | F7 | 4
67 | F7 | 5
67 | F7 | 6
67 | F7 | 7
67 | F7 | 8
67 | F7 | 9
67 | F7 | 10
67 | F7 | 11
67 | F7 | 12
67 | F7 | 13
67 | F7 | 14
67 | F7 | 15
67 | F7 | 16
67 | F7 | 17
67 | F7 | 18
67 | F7 | 19
67 | F7 | 20
67 | F7 | 21
67 | F7 | 22
67 | F7 | 23
67 | F7 | 24
67 | F7 | 25
67 | F7 | 26
67 | F7 | 27
67 | F7 | 28
67 | F7 | 29
67 | F7 | 30
67 | F7 | 31
67 | F7 | 32
67 | F7 | 33
67 | F7 | 34
67 | F7 | 35
68 | F8 | 1
68 | F8 | 2
68 | F8 | 3
68 | F8 | 4
68 | F8 | 5
68 | F8 | 6
68 | F8 | 7
68 | F8 | 8
68 | F8 | 9
68 | F8 | 10
68 | F8 | 11
68 | F8 | 12
68 | F8 | 13
68 | F8 | 14
68 | F8 | 15
68 | F8 | 16
68 | F8 | 17
68 | F8 | 18
68 | F8 | 19
68 | F8 | 20
68 | F8 | 21
68 | F8 | 22
68 | F8 | 23
68 | F8 | 24
68 | F8 | 25
68 | F8 | 26
68 | F8 | 27
68 | F8 | 28
68 | F8 | 29
68 | F8 | 30
68 | F8 | 31
68 | F8 | 32
68 | F8 | 33
68 | F8 | 34
68 | F8 | 35
69 | F9 | 1
69 | F9 | 2
69 | F9 | 3
69 | F9 | 4
69 | F9 | 5
69 | F9 | 6
69 | F9 | 7
69 | F9 | 8
69 | F9 | 9
69 | F9 | 10
69 | F9 | 11
69 | F9 | 12
69 | F9 | 13
69 | F9 | 14
69 | F9 | 15
69 | F9 | 16
69 | F9 | 17
69 | F9 | 18
69 | F9 | 19
69 | F9 | 20
69 | F9 | 21
69 | F9 | 22
69 | F9 | 23
69 | F9 | 24
69 | F9 | 25
69 | F9 | 26
69 | F9 | 27
69 | F9 | 28
69 | F9 | 29
69 | F9 | 30
69 | F9 | 31
69 | F9 | 32
69 | F9 | 33
69 | F9 | 34
69 | F9 | 35
70 | F10 | 1
70 | F10 | 2
70 | F10 | 3
70 | F10 | 4
70 | F10 | 5
70 | F10 | 6
70 | F10 | 7
70 | F10 | 8
70 | F10 | 9
70 | F10 | 10
70 | F10 | 11
70 | F10 | 12
70 | F10 | 13
70 | F10 | 14
70 | F10 | 15
70 | F10 | 16
70 | F10 | 17
70 | F10 | 18
70 | F10 | 19
70 | F10 | 20
70 | F10 | 21
70 | F10 | 22
70 | F10 | 23
70 | F10 | 24
70 | F10 | 25
70 | F10 | 26
70 | F10 | 27
70 | F10 | 28
70 | F10 | 29
70 | F10 | 30
70 | F10 | 31
70 | F10 | 32
70 | F10 | 33
70 | F10 | 34
70 | F10 | 35
71 | F11 | 1
71 | F11 | 2
71 | F11 | 3
71 | F11 | 4
71 | F11 | 5
71 | F11 | 6
71 | F11 | 7
71 | F11 | 8
71 | F11 | 9
71 | F11 | 10
71 | F11 | 11
71 | F11 | 12
71 | F11 | 13
71 | F11 | 14
71 | F11 | 15
71 | F11 | 16
71 | F11 | 17
71 | F11 | 18
71 | F11 | 19
71 | F11 | 20
71 | F11 | 21
71 | F11 | 22
71 | F11 | 23
71 | F11 | 24
71 | F11 | 25
71 | F11 | 26
71 | F11 | 27
71 | F11 | 28
71 | F11 | 29
71 | F11 | 30
71 | F11 | 31
71 | F11 | 32
71 | F11 | 33
71 | F11 | 34
71 | F11 | 35
72 | F12 | 1
72 | F12 | 2
72 | F12 | 3
72 | F12 | 4
72 | F12 | 5
72 | F12 | 6
72 | F12 | 7
72 | F12 | 8
72 | F12 | 9
72 | F12 | 10
72 | F12 | 11
72 | F12 | 12
72 | F12 | 13
72 | F12 | 14
72 | F12 | 15
72 | F12 | 16
72 | F12 | 17
72 | F12 | 18
72 | F12 | 19
72 | F12 | 20
72 | F12 | 21
72 | F12 | 22
72 | F12 | 23
72 | F12 | 24
72 | F12 | 25
72 | F12 | 26
72 | F12 | 27
72 | F12 | 28
72 | F12 | 29
72 | F12 | 30
72 | F12 | 31
72 | F12 | 32
72 | F12 | 33
72 | F12 | 34
72 | F12 | 35
73 | G1 | 1
73 | G1 | 2
73 | G1 | 3
73 | G1 | 4
73 | G1 | 5
73 | G1 | 6
73 | G1 | 7
73 | G1 | 8
73 | G1 | 9
73 | G1 | 10
73 | G1 | 11
73 | G1 | 12
73 | G1 | 13
73 | G1 | 14
73 | G1 | 15
73 | G1 | 16
73 | G1 | 17
73 | G1 | 18
73 | G1 | 19
73 | G1 | 20
73 | G1 | 21
73 | G1 | 22
73 | G1 | 23
73 | G1 | 24
73 | G1 | 25
73 | G1 | 26
73 | G1 | 27
73 | G1 | 28
73 | G1 | 29
73 | G1 | 30
73 | G1 | 31
73 | G1 | 32
73 | G1 | 33
73 | G1 | 34
73 | G1 | 35
74 | G2 | 1
74 | G2 | 2
74 | G2 | 3
74 | G2 | 4
74 | G2 | 5
74 | G2 | 6
74 | G2 | 7
74 | G2 | 8
74 | G2 | 9
74 | G2 | 10
74 | G2 | 11
74 | G2 | 12
74 | G2 | 13
74 | G2 | 14
74 | G2 | 15
74 | G2 | 16
74 | G2 | 17
74 | G2 | 18
74 | G2 | 19
74 | G2 | 20
74 | G2 | 21
74 | G2 | 22
74 | G2 | 23
74 | G2 | 24
74 | G2 | 25
74 | G2 | 26
74 | G2 | 27
74 | G2 | 28
74 | G2 | 29
74 | G2 | 30
74 | G2 | 31
74 | G2 | 32
74 | G2 | 33
74 | G2 | 34
74 | G2 | 35
75 | G3 | 1 | RNP | 0.084388 | 0.0031331
75 | G3 | 2 | RNP | 0.083681 | 0.0022444
75 | G3 | 3 | RNP | 0.083053 | 0.0014349
75 | G3 | 4 | RNP | 0.084427 | 0.0026267
75 | G3 | 5 | RNP | 0.084029 | 0.0020469
75 | G3 | 6 | RNP | 0.083527 | 0.0013638
75 | G3 | 7 | RNP | 0.083655 | 0.0013094
75 | G3 | 8 | RNP | 0.081829 | -0.00069845
75 | G3 | 9 | RNP | 0.08179 | -0.00091957
75 | G3 | 10 | RNP | 0.084676 | 0.001785
75 | G3 | 11 | RNP | 0.082928 | -0.00014509
75 | G3 | 12 | RNP | 0.082171 | -0.001084
75 | G3 | 13 | RNP | 0.081745 | -0.0016913
75 | G3 | 14 | RNP | 0.079547 | -0.004071
75 | G3 | 15 | RNP | 0.07987 | -0.0039305
75 | G3 | 16 | RNP | 0.082832 | -0.0011502
75 | G3 | 17 | RNP | 0.082953 | -0.0012104
75 | G3 | 18 | RNP | 0.080817 | -0.0035291
75 | G3 | 19 | RNP | 0.080734 | -0.0037935
75 | G3 | 20 | RNP | 0.082868 | -0.0018417
75 | G3 | 21 | RNP | 0.083927 | -0.00096391
75 | G3 | 22 | RNP | 0.087299 | 0.0022256
75 | G3 | 23 | RNP | 0.090075 | 0.0048203
75 | G3 | 24 | RNP | 0.092853 | 0.0074162
75 | G3 | 25 | RNP | 0.09948 | 0.013861
75 | G3 | 26 | RNP | 0.11593 | 0.030132
75 | G3 | 27 | RNP | 0.14069 | 0.05471
75 | G3 | 28 | RNP | 0.17429 | 0.088126
75 | G3 | 29 | RNP | 0.21346 | 0.12711
75 | G3 | 30 | RNP | 0.25725 | 0.17072
75 | G3 | 31 | RNP | 0.30026 | 0.21355
75 | G3 | 32 | RNP | 0.3397 | 0.25281
75 | G3 | 33 | RNP | 0.37746 | 0.29039
75 | G3 | 34 | RNP | 0.40889 | 0.32163
75 | G3 | 35 | RNP | 0.43503 | 0.34759
75 | G3 | 1 | SMN2 | 0.33404 | 0.002057
75 | G3 | 2 | SMN2 | 0.33354 | 0.0015504
75 | G3 | 3 | SMN2 | 0.33239 | 0.0003984
75 | G3 | 4 | SMN2 | 0.33504 | 0.00305
75 | G3 | 5 | SMN2 | 0.33478 | 0.0027855
75 | G3 | 6 | SMN2 | 0.33298 | 0.0009737
75 | G3 | 7 | SMN2 | 0.33206 | 5.5773e-05
75 | G3 | 8 | SMN2 | 0.33164 | -0.00037053
75 | G3 | 9 | SMN2 | 0.33176 | -0.00025837
75 | G3 | 10 | SMN2 | 0.33326 | 0.0012399
75 | G3 | 11 | SMN2 | 0.33199 | -3.8151e-05
75 | G3 | 12 | SMN2 | 0.33004 | -0.0019915
75 | G3 | 13 | SMN2 | 0.32759 | -0.004442
75 | G3 | 14 | SMN2 | 0.32803 | -0.0040052
75 | G3 | 15 | SMN2 | 0.3307 | -0.0013461
75 | G3 | 16 | SMN2 | 0.33051 | -0.0015321
75 | G3 | 17 | SMN2 | 0.33045 | -0.0016019
75 | G3 | 18 | SMN2 | 0.3305 | -0.0015586
75 | G3 | 19 | SMN2 | 0.33068 | -0.0013808
75 | G3 | 20 | SMN2 | 0.33154 | -0.00052767
75 | G3 | 21 | SMN2 | 0.33445 | 0.0023816
75 | G3 | 22 | SMN2 | 0.3357 | 0.0036264
75 | G3 | 23 | SMN2 | 0.33662 | 0.0045417
75 | G3 | 24 | SMN2 | 0.34402 | 0.011935
75 | G3 | 25 | SMN2 | 0.35889 | 0.026803
75 | G3 | 26 | SMN2 | 0.38368 | 0.051588
75 | G3 | 27 | SMN2 | 0.42985 | 0.097753
75 | G3 | 28 | SMN2 | 0.49526 | 0.16316
75 | G3 | 29 | SMN2 | 0.57231 | 0.24021
75 | G3 | 30 | SMN2 | 0.65514 | 0.32304
75 | G3 | 31 | SMN2 | 0.72848 | 0.39637
75 | G3 | 32 | SMN2 | 0.79616 | 0.46405
75 | G3 | 33 | SMN2 | 0.86028 | 0.52816
75 | G3 | 34 | SMN2 | 0.92585 | 0.59373
75 | G3 | 35 | SMN2 | 0.9758 | 0.64368
75 | G3 | 1 | SMN1 | 0.48092 | 0.003508
75 | G3 | 2 | SMN1 | 0.48001 | 0.0029333
75 | G3 | 3 | SMN1 | 0.47933 | 0.002583
75 | G3 | 4 | SMN1 | 0.48105 | 0.0046443
75 | G3 | 5 | SMN1 | 0.48039 | 0.004313
75 | G3 | 6 | SMN1 | 0.47509 | -0.00064526
75 | G3 | 7 | SMN1 | 0.4738 | -0.0015977
75 | G3 | 8 | SMN1 | 0.47436 | -0.00070075
75 | G3 | 9 | SMN1 | 0.46979 | -0.0049355
75 | G3 | 10 | SMN1 | 0.46892 | -0.005471
75 | G3 | 11 | SMN1 | 0.47021 | -0.0038434
75 | G3 | 12 | SMN1 | 0.47332 | -0.0004013
75 | G3 | 13 | SMN1 | 0.47428 | 0.00089701
75 | G3 | 14 | SMN1 | 0.46985 | -0.0031956
75 | G3 | 15 | SMN1 | 0.47049 | -0.0022286
75 | G3 | 16 | SMN1 | 0.47429 | 0.0019139
75 | G3 | 17 | SMN1 | 0.47546 | 0.0034172
75 | G3 | 18 | SMN1 | 0.47469 | 0.0029798
75 | G3 | 19 | SMN1 | 0.47364 | 0.002271
75 | G3 | 20 | SMN1 | 0.47572 | 0.0046831
75 | G3 | 21 | SMN1 | 0.48644 | 0.01574
75 | G3 | 22 | SMN1 | 0.49939 | 0.02903
75 | G3 | 23 | SMN1 | 0.51536 | 0.045336
75 | G3 | 24 | SMN1 | 0.54899 | 0.079296
75 | G3 | 25 | SMN1 | 0.62377 | 0.15441
75 | G3 | 26 | SMN1 | 0.74197 | 0.27295
75 | G3 | 27 | SMN1 | 0.88167 | 0.41298
75 | G3 | 28 | SMN1 | 1.0303 | 0.56192
75 | G3 | 29 | SMN1 | 1.1752 | 0.70718
75 | G3 | 30 | SMN1 | 1.3115 | 0.84385
75 | G3 | 31 | SMN1 | 1.4384 | 0.97105
75 | G3 | 32 | SMN1 | 1.5553 | 1.0883
75 | G3 | 33 | SMN1 | 1.6808 | 1.2141
75 | G3 | 34 | SMN1 | 1.8042 | 1.3378
75 | G3 | 35 | SMN1 | 1.8996 | 1.4336
76 | G4 | 1
76 | G4 | 2
76 | G4 | 3
76 | G4 | 4
76 | G4 | 5
76 | G4 | 6
76 | G4 | 7
76 | G4 | 8
76 | G4 | 9
76 | G4 | 10
76 | G4 | 11
76 | G4 | 12
76 | G4 | 13
76 | G4 | 14
76 | G4 | 15
76 | G4 | 16
76 | G4 | 17
76 | G4 | 18
76 | G4 | 19
76 | G4 | 20
76 | G4 | 21
76 | G4 | 22
76 | G4 | 23
76 | G4 | 24
76 | G4 | 25
76 | G4 | 26
76 | G4 | 27
76 | G4 | 28
76 | G4 | 29
76 | G4 | 30
76 | G4 | 31
76 | G4 | 32
76 | G4 | 33
76 | G4 | 34
76 | G4 | 35
77 | G5 | 1
77 | G5 | 2
77 | G5 | 3
77 | G5 | 4
77 | G5 | 5
77 | G5 | 6
77 | G5 | 7
77 | G5 | 8
77 | G5 | 9
77 | G5 | 10
77 | G5 | 11
77 | G5 | 12
77 | G5 | 13
77 | G5 | 14
77 | G5 | 15
77 | G5 | 16
77 | G5 | 17
77 | G5 | 18
77 | G5 | 19
77 | G5 | 20
77 | G5 | 21
77 | G5 | 22
77 | G5 | 23
77 | G5 | 24
77 | G5 | 25
77 | G5 | 26
77 | G5 | 27
77 | G5 | 28
77 | G5 | 29
77 | G5 | 30
77 | G5 | 31
77 | G5 | 32
77 | G5 | 33
77 | G5 | 34
77 | G5 | 35
78 | G6 | 1
78 | G6 | 2
78 | G6 | 3
78 | G6 | 4
78 | G6 | 5
78 | G6 | 6
78 | G6 | 7
78 | G6 | 8
78 | G6 | 9
78 | G6 | 10
78 | G6 | 11
78 | G6 | 12
78 | G6 | 13
78 | G6 | 14
78 | G6 | 15
78 | G6 | 16
78 | G6 | 17
78 | G6 | 18
78 | G6 | 19
78 | G6 | 20
78 | G6 | 21
78 | G6 | 22
78 | G6 | 23
78 | G6 | 24
78 | G6 | 25
78 | G6 | 26
78 | G6 | 27
78 | G6 | 28
78 | G6 | 29
78 | G6 | 30
78 | G6 | 31
78 | G6 | 32
78 | G6 | 33
78 | G6 | 34
78 | G6 | 35
79 | G7 | 1
79 | G7 | 2
79 | G7 | 3
79 | G7 | 4
79 | G7 | 5
79 | G7 | 6
79 | G7 | 7
79 | G7 | 8
79 | G7 | 9
79 | G7 | 10
79 | G7 | 11
79 | G7 | 12
79 | G7 | 13
79 | G7 | 14
79 | G7 | 15
79 | G7 | 16
79 | G7 | 17
79 | G7 | 18
79 | G7 | 19
79 | G7 | 20
79 | G7 | 21
79 | G7 | 22
79 | G7 | 23
79 | G7 | 24
79 | G7 | 25
79 | G7 | 26
79 | G7 | 27
79 | G7 | 28
79 | G7 | 29
79 | G7 | 30
79 | G7 | 31
79 | G7 | 32
79 | G7 | 33
79 | G7 | 34
79 | G7 | 35
80 | G8 | 1
80 | G8 | 2
80 | G8 | 3
80 | G8 | 4
80 | G8 | 5
80 | G8 | 6
80 | G8 | 7
80 | G8 | 8
80 | G8 | 9
80 | G8 | 10
80 | G8 | 11
80 | G8 | 12
80 | G8 | 13
80 | G8 | 14
80 | G8 | 15
80 | G8 | 16
80 | G8 | 17
80 | G8 | 18
80 | G8 | 19
80 | G8 | 20
80 | G8 | 21
80 | G8 | 22
80 | G8 | 23
80 | G8 | 24
80 | G8 | 25
80 | G8 | 26
80 | G8 | 27
80 | G8 | 28
80 | G8 | 29
80 | G8 | 30
80 | G8 | 31
80 | G8 | 32
80 | G8 | 33
80 | G8 | 34
80 | G8 | 35
81 | G9 | 1
81 | G9 | 2
81 | G9 | 3
81 | G9 | 4
81 | G9 | 5
81 | G9 | 6
81 | G9 | 7
81 | G9 | 8
81 | G9 | 9
81 | G9 | 10
81 | G9 | 11
81 | G9 | 12
81 | G9 | 13
81 | G9 | 14
81 | G9 | 15
81 | G9 | 16
81 | G9 | 17
81 | G9 | 18
81 | G9 | 19
81 | G9 | 20
81 | G9 | 21
81 | G9 | 22
81 | G9 | 23
81 | G9 | 24
81 | G9 | 25
81 | G9 | 26
81 | G9 | 27
81 | G9 | 28
81 | G9 | 29
81 | G9 | 30
81 | G9 | 31
81 | G9 | 32
81 | G9 | 33
81 | G9 | 34
81 | G9 | 35
82 | G10 | 1
82 | G10 | 2
82 | G10 | 3
82 | G10 | 4
82 | G10 | 5
82 | G10 | 6
82 | G10 | 7
82 | G10 | 8
82 | G10 | 9
82 | G10 | 10
82 | G10 | 11
82 | G10 | 12
82 | G10 | 13
82 | G10 | 14
82 | G10 | 15
82 | G10 | 16
82 | G10 | 17
82 | G10 | 18
82 | G10 | 19
82 | G10 | 20
82 | G10 | 21
82 | G10 | 22
82 | G10 | 23
82 | G10 | 24
82 | G10 | 25
82 | G10 | 26
82 | G10 | 27
82 | G10 | 28
82 | G10 | 29
82 | G10 | 30
82 | G10 | 31
82 | G10 | 32
82 | G10 | 33
82 | G10 | 34
82 | G10 | 35
83 | G11 | 1
83 | G11 | 2
83 | G11 | 3
83 | G11 | 4
83 | G11 | 5
83 | G11 | 6
83 | G11 | 7
83 | G11 | 8
83 | G11 | 9
83 | G11 | 10
83 | G11 | 11
83 | G11 | 12
83 | G11 | 13
83 | G11 | 14
83 | G11 | 15
83 | G11 | 16
83 | G11 | 17
83 | G11 | 18
83 | G11 | 19
83 | G11 | 20
83 | G11 | 21
83 | G11 | 22
83 | G11 | 23
83 | G11 | 24
83 | G11 | 25
83 | G11 | 26
83 | G11 | 27
83 | G11 | 28
83 | G11 | 29
83 | G11 | 30
83 | G11 | 31
83 | G11 | 32
83 | G11 | 33
83 | G11 | 34
83 | G11 | 35
84 | G12 | 1
84 | G12 | 2
84 | G12 | 3
84 | G12 | 4
84 | G12 | 5
84 | G12 | 6
84 | G12 | 7
84 | G12 | 8
84 | G12 | 9
84 | G12 | 10
84 | G12 | 11
84 | G12 | 12
84 | G12 | 13
84 | G12 | 14
84 | G12 | 15
84 | G12 | 16
84 | G12 | 17
84 | G12 | 18
84 | G12 | 19
84 | G12 | 20
84 | G12 | 21
84 | G12 | 22
84 | G12 | 23
84 | G12 | 24
84 | G12 | 25
84 | G12 | 26
84 | G12 | 27
84 | G12 | 28
84 | G12 | 29
84 | G12 | 30
84 | G12 | 31
84 | G12 | 32
84 | G12 | 33
84 | G12 | 34
84 | G12 | 35
85 | H1 | 1
85 | H1 | 2
85 | H1 | 3
85 | H1 | 4
85 | H1 | 5
85 | H1 | 6
85 | H1 | 7
85 | H1 | 8
85 | H1 | 9
85 | H1 | 10
85 | H1 | 11
85 | H1 | 12
85 | H1 | 13
85 | H1 | 14
85 | H1 | 15
85 | H1 | 16
85 | H1 | 17
85 | H1 | 18
85 | H1 | 19
85 | H1 | 20
85 | H1 | 21
85 | H1 | 22
85 | H1 | 23
85 | H1 | 24
85 | H1 | 25
85 | H1 | 26
85 | H1 | 27
85 | H1 | 28
85 | H1 | 29
85 | H1 | 30
85 | H1 | 31
85 | H1 | 32
85 | H1 | 33
85 | H1 | 34
85 | H1 | 35
86 | H2 | 1
86 | H2 | 2
86 | H2 | 3
86 | H2 | 4
86 | H2 | 5
86 | H2 | 6
86 | H2 | 7
86 | H2 | 8
86 | H2 | 9
86 | H2 | 10
86 | H2 | 11
86 | H2 | 12
86 | H2 | 13
86 | H2 | 14
86 | H2 | 15
86 | H2 | 16
86 | H2 | 17
86 | H2 | 18
86 | H2 | 19
86 | H2 | 20
86 | H2 | 21
86 | H2 | 22
86 | H2 | 23
86 | H2 | 24
86 | H2 | 25
86 | H2 | 26
86 | H2 | 27
86 | H2 | 28
86 | H2 | 29
86 | H2 | 30
86 | H2 | 31
86 | H2 | 32
86 | H2 | 33
86 | H2 | 34
86 | H2 | 35
87 | H3 | 1 | RNP | 0.091541 | 0.00010039
87 | H3 | 2 | RNP | 0.092708 | 0.001384
87 | H3 | 3 | RNP | 0.092903 | 0.0016949
87 | H3 | 4 | RNP | 0.093055 | 0.0019627
87 | H3 | 5 | RNP | 0.091925 | 0.0009489
87 | H3 | 6 | RNP | 0.091127 | 0.00026786
87 | H3 | 7 | RNP | 0.090558 | -0.00018524
87 | H3 | 8 | RNP | 0.090581 | -4.6243e-05
87 | H3 | 9 | RNP | 0.091694 | 0.0011833
87 | H3 | 10 | RNP | 0.089151 | -0.0012434
87 | H3 | 11 | RNP | 0.088384 | -0.0018946
87 | H3 | 12 | RNP | 0.08921 | -0.00095219
87 | H3 | 13 | RNP | 0.089228 | -0.00081832
87 | H3 | 14 | RNP | 0.088756 | -0.0011738
87 | H3 | 15 | RNP | 0.086284 | -0.0035295
87 | H3 | 16 | RNP | 0.086336 | -0.0033619
87 | H3 | 17 | RNP | 0.088681 | -0.00090069
87 | H3 | 18 | RNP | 0.089381 | -8.4617e-05
87 | H3 | 19 | RNP | 0.091217 | 0.0018682
87 | H3 | 20 | RNP | 0.091403 | 0.0021705
87 | H3 | 21 | RNP | 0.091025 | 0.0019081
87 | H3 | 22 | RNP | 0.091187 | 0.0021861
87 | H3 | 23 | RNP | 0.095728 | 0.0068442
87 | H3 | 24 | RNP | 0.10473 | 0.015964
87 | H3 | 25 | RNP | 0.11672 | 0.028072
87 | H3 | 26 | RNP | 0.13952 | 0.050982
87 | H3 | 27 | RNP | 0.17107 | 0.08265
87 | H3 | 28 | RNP | 0.21338 | 0.12508
87 | H3 | 29 | RNP | 0.25849 | 0.1703
87 | H3 | 30 | RNP | 0.30542 | 0.21735
87 | H3 | 31 | RNP | 0.34859 | 0.26063
87 | H3 | 32 | RNP | 0.38122 | 0.29338
87 | H3 | 33 | RNP | 0.41357 | 0.32584
87 | H3 | 34 | RNP | 0.44183 | 0.35423
87 | H3 | 35 | RNP | 0.46125 | 0.37376
87 | H3 | 1 | SMN2 | 0.35209 | 0.0032383
87 | H3 | 2 | SMN2 | 0.352 | 0.0030217
87 | H3 | 3 | SMN2 | 0.35094 | 0.001843
87 | H3 | 4 | SMN2 | 0.3507 | 0.0014868
87 | H3 | 5 | SMN2 | 0.3511 | 0.0017657
87 | H3 | 6 | SMN2 | 0.34973 | 0.00026961
87 | H3 | 7 | SMN2 | 0.34887 | -0.00070387
87 | H3 | 8 | SMN2 | 0.3476 | -0.0021026
87 | H3 | 9 | SMN2 | 0.3477 | -0.0021235
87 | H3 | 10 | SMN2 | 0.34872 | -0.0012232
87 | H3 | 11 | SMN2 | 0.34916 | -0.00090495
87 | H3 | 12 | SMN2 | 0.34969 | -0.00048646
87 | H3 | 13 | SMN2 | 0.34894 | -0.0013592
87 | H3 | 14 | SMN2 | 0.34984 | -0.00057831
87 | H3 | 15 | SMN2 | 0.35201 | 0.0014705
87 | H3 | 16 | SMN2 | 0.35273 | 0.0020671
87 | H3 | 17 | SMN2 | 0.35044 | -0.00033956
87 | H3 | 18 | SMN2 | 0.3496 | -0.0013072
87 | H3 | 19 | SMN2 | 0.3505 | -0.00052374
87 | H3 | 20 | SMN2 | 0.35204 | 0.00089752
87 | H3 | 21 | SMN2 | 0.35312 | 0.0018523
87 | H3 | 22 | SMN2 | 0.35745 | 0.0060624
87 | H3 | 23 | SMN2 | 0.36505 | 0.013545
87 | H3 | 24 | SMN2 | 0.3788 | 0.027175
87 | H3 | 25 | SMN2 | 0.41854 | 0.066787
87 | H3 | 26 | SMN2 | 0.48309 | 0.13122
87 | H3 | 27 | SMN2 | 0.58398 | 0.23199
87 | H3 | 28 | SMN2 | 0.70108 | 0.34896
87 | H3 | 29 | SMN2 | 0.82798 | 0.47575
87 | H3 | 30 | SMN2 | 0.94318 | 0.59083
87 | H3 | 31 | SMN2 | 1.0467 | 0.69426
87 | H3 | 32 | SMN2 | 1.1403 | 0.78775
87 | H3 | 33 | SMN2 | 1.2347 | 0.88194
87 | H3 | 34 | SMN2 | 1.3145 | 0.96169
87 | H3 | 35 | SMN2 | 1.3774 | 1.0244
87 | H3 | 1 | SMN1 | 0.49864 | 0.0036236
87 | H3 | 2 | SMN1 | 0.4995 | 0.0044447
87 | H3 | 3 | SMN1 | 0.49965 | 0.0045615
87 | H3 | 4 | SMN1 | 0.49968 | 0.0045576
87 | H3 | 5 | SMN1 | 0.50024 | 0.0050771
87 | H3 | 6 | SMN1 | 0.49764 | 0.0024497
87 | H3 | 7 | SMN1 | 0.48883 | -0.0064025
87 | H3 | 8 | SMN1 | 0.48838 | -0.0068798
87 | H3 | 9 | SMN1 | 0.49451 | -0.00078632
87 | H3 | 10 | SMN1 | 0.49575 | 0.00041427
87 | H3 | 11 | SMN1 | 0.49393 | -0.0014426
87 | H3 | 12 | SMN1 | 0.4931 | -0.0023011
87 | H3 | 13 | SMN1 | 0.49228 | -0.0031576
87 | H3 | 14 | SMN1 | 0.49352 | -0.0019573
87 | H3 | 15 | SMN1 | 0.4956 | 8.5283e-05
87 | H3 | 16 | SMN1 | 0.49509 | -0.00045804
87 | H3 | 17 | SMN1 | 0.49191 | -0.003673
87 | H3 | 18 | SMN1 | 0.4922 | -0.0034148
87 | H3 | 19 | SMN1 | 0.49732 | 0.001665
87 | H3 | 20 | SMN1 | 0.49956 | 0.0038687
87 | H3 | 21 | SMN1 | 0.50352 | 0.0077938
87 | H3 | 22 | SMN1 | 0.50795 | 0.012188
87 | H3 | 23 | SMN1 | 0.52116 | 0.02536
87 | H3 | 24 | SMN1 | 0.55245 | 0.056622
87 | H3 | 25 | SMN1 | 0.60085 | 0.10498
87 | H3 | 26 | SMN1 | 0.66686 | 0.17096
87 | H3 | 27 | SMN1 | 0.74398 | 0.24804
87 | H3 | 28 | SMN1 | 0.81755 | 0.32158
87 | H3 | 29 | SMN1 | 0.88958 | 0.39358
87 | H3 | 30 | SMN1 | 0.95801 | 0.46197
87 | H3 | 31 | SMN1 | 1.0227 | 0.52659
87 | H3 | 32 | SMN1 | 1.0815 | 0.58539
87 | H3 | 33 | SMN1 | 1.1498 | 0.65363
87 | H3 | 34 | SMN1 | 1.2114 | 0.7152
87 | H3 | 35 | SMN1 | 1.2609 | 0.76466
88 | H4 | 1
88 | H4 | 2
88 | H4 | 3
88 | H4 | 4
88 | H4 | 5
88 | H4 | 6
88 | H4 | 7
88 | H4 | 8
88 | H4 | 9
88 | H4 | 10
88 | H4 | 11
88 | H4 | 12
88 | H4 | 13
88 | H4 | 14
88 | H4 | 15
88 | H4 | 16
88 | H4 | 17
88 | H4 | 18
88 | H4 | 19
88 | H4 | 20
88 | H4 | 21
88 | H4 | 22
88 | H4 | 23
88 | H4 | 24
88 | H4 | 25
88 | H4 | 26
88 | H4 | 27
88 | H4 | 28
88 | H4 | 29
88 | H4 | 30
88 | H4 | 31
88 | H4 | 32
88 | H4 | 33
88 | H4 | 34
88 | H4 | 35
89 | H5 | 1
89 | H5 | 2
89 | H5 | 3
89 | H5 | 4
89 | H5 | 5
89 | H5 | 6
89 | H5 | 7
89 | H5 | 8
89 | H5 | 9
89 | H5 | 10
89 | H5 | 11
89 | H5 | 12
89 | H5 | 13
89 | H5 | 14
89 | H5 | 15
89 | H5 | 16
89 | H5 | 17
89 | H5 | 18
89 | H5 | 19
89 | H5 | 20
89 | H5 | 21
89 | H5 | 22
89 | H5 | 23
89 | H5 | 24
89 | H5 | 25
89 | H5 | 26
89 | H5 | 27
89 | H5 | 28
89 | H5 | 29
89 | H5 | 30
89 | H5 | 31
89 | H5 | 32
89 | H5 | 33
89 | H5 | 34
89 | H5 | 35
90 | H6 | 1
90 | H6 | 2
90 | H6 | 3
90 | H6 | 4
90 | H6 | 5
90 | H6 | 6
90 | H6 | 7
90 | H6 | 8
90 | H6 | 9
90 | H6 | 10
90 | H6 | 11
90 | H6 | 12
90 | H6 | 13
90 | H6 | 14
90 | H6 | 15
90 | H6 | 16
90 | H6 | 17
90 | H6 | 18
90 | H6 | 19
90 | H6 | 20
90 | H6 | 21
90 | H6 | 22
90 | H6 | 23
90 | H6 | 24
90 | H6 | 25
90 | H6 | 26
90 | H6 | 27
90 | H6 | 28
90 | H6 | 29
90 | H6 | 30
90 | H6 | 31
90 | H6 | 32
90 | H6 | 33
90 | H6 | 34
90 | H6 | 35
91 | H7 | 1
91 | H7 | 2
91 | H7 | 3
91 | H7 | 4
91 | H7 | 5
91 | H7 | 6
91 | H7 | 7
91 | H7 | 8
91 | H7 | 9
91 | H7 | 10
91 | H7 | 11
91 | H7 | 12
91 | H7 | 13
91 | H7 | 14
91 | H7 | 15
91 | H7 | 16
91 | H7 | 17
91 | H7 | 18
91 | H7 | 19
91 | H7 | 20
91 | H7 | 21
91 | H7 | 22
91 | H7 | 23
91 | H7 | 24
91 | H7 | 25
91 | H7 | 26
91 | H7 | 27
91 | H7 | 28
91 | H7 | 29
91 | H7 | 30
91 | H7 | 31
91 | H7 | 32
91 | H7 | 33
91 | H7 | 34
91 | H7 | 35
92 | H8 | 1
92 | H8 | 2
92 | H8 | 3
92 | H8 | 4
92 | H8 | 5
92 | H8 | 6
92 | H8 | 7
92 | H8 | 8
92 | H8 | 9
92 | H8 | 10
92 | H8 | 11
92 | H8 | 12
92 | H8 | 13
92 | H8 | 14
92 | H8 | 15
92 | H8 | 16
92 | H8 | 17
92 | H8 | 18
92 | H8 | 19
92 | H8 | 20
92 | H8 | 21
92 | H8 | 22
92 | H8 | 23
92 | H8 | 24
92 | H8 | 25
92 | H8 | 26
92 | H8 | 27
92 | H8 | 28
92 | H8 | 29
92 | H8 | 30
92 | H8 | 31
92 | H8 | 32
92 | H8 | 33
92 | H8 | 34
92 | H8 | 35
93 | H9 | 1
93 | H9 | 2
93 | H9 | 3
93 | H9 | 4
93 | H9 | 5
93 | H9 | 6
93 | H9 | 7
93 | H9 | 8
93 | H9 | 9
93 | H9 | 10
93 | H9 | 11
93 | H9 | 12
93 | H9 | 13
93 | H9 | 14
93 | H9 | 15
93 | H9 | 16
93 | H9 | 17
93 | H9 | 18
93 | H9 | 19
93 | H9 | 20
93 | H9 | 21
93 | H9 | 22
93 | H9 | 23
93 | H9 | 24
93 | H9 | 25
93 | H9 | 26
93 | H9 | 27
93 | H9 | 28
93 | H9 | 29
93 | H9 | 30
93 | H9 | 31
93 | H9 | 32
93 | H9 | 33
93 | H9 | 34
93 | H9 | 35
94 | H10 | 1
94 | H10 | 2
94 | H10 | 3
94 | H10 | 4
94 | H10 | 5
94 | H10 | 6
94 | H10 | 7
94 | H10 | 8
94 | H10 | 9
94 | H10 | 10
94 | H10 | 11
94 | H10 | 12
94 | H10 | 13
94 | H10 | 14
94 | H10 | 15
94 | H10 | 16
94 | H10 | 17
94 | H10 | 18
94 | H10 | 19
94 | H10 | 20
94 | H10 | 21
94 | H10 | 22
94 | H10 | 23
94 | H10 | 24
94 | H10 | 25
94 | H10 | 26
94 | H10 | 27
94 | H10 | 28
94 | H10 | 29
94 | H10 | 30
94 | H10 | 31
94 | H10 | 32
94 | H10 | 33
94 | H10 | 34
94 | H10 | 35
95 | H11 | 1
95 | H11 | 2
95 | H11 | 3
95 | H11 | 4
95 | H11 | 5
95 | H11 | 6
95 | H11 | 7
95 | H11 | 8
95 | H11 | 9
95 | H11 | 10
95 | H11 | 11
95 | H11 | 12
95 | H11 | 13
95 | H11 | 14
95 | H11 | 15
95 | H11 | 16
95 | H11 | 17
95 | H11 | 18
95 | H11 | 19
95 | H11 | 20
95 | H11 | 21
95 | H11 | 22
95 | H11 | 23
95 | H11 | 24
95 | H11 | 25
95 | H11 | 26
95 | H11 | 27
95 | H11 | 28
95 | H11 | 29
95 | H11 | 30
95 | H11 | 31
95 | H11 | 32
95 | H11 | 33
95 | H11 | 34
95 | H11 | 35
96 | H12 | 1
96 | H12 | 2
96 | H12 | 3
96 | H12 | 4
96 | H12 | 5
96 | H12 | 6
96 | H12 | 7
96 | H12 | 8
96 | H12 | 9
96 | H12 | 10
96 | H12 | 11
96 | H12 | 12
96 | H12 | 13
96 | H12 | 14
96 | H12 | 15
96 | H12 | 16
96 | H12 | 17
96 | H12 | 18
96 | H12 | 19
96 | H12 | 20
96 | H12 | 21
96 | H12 | 22
96 | H12 | 23
96 | H12 | 24
96 | H12 | 25
96 | H12 | 26
96 | H12 | 27
96 | H12 | 28
96 | H12 | 29
96 | H12 | 30
96 | H12 | 31
96 | H12 | 32
96 | H12 | 33
96 | H12 | 34
96 | H12 | 35

## Results
Block Type | 96-Well 0.2-mL Block
Calibration Background is expired | No
Calibration Background performed on | 06-01-2021
Calibration Pure Dye ABY is expired | No
Calibration Pure Dye ABY performed on | 06-01-2021
Calibration Pure Dye FAM is expired | No
Calibration Pure Dye FAM performed on | 06-01-2021
Calibration Pure Dye JUN is expired | No
Calibration Pure Dye JUN performed on | 06-01-2021
Calibration Pure Dye NED is expired | No
Calibration Pure Dye NED performed on | 06-01-2021
Calibration Pure Dye ROX is expired | No
Calibration Pure Dye ROX performed on | 06-01-2021
Calibration Pure Dye SYBR is expired | No
Calibration Pure Dye SYBR performed on | 06-01-2021
Calibration Pure Dye TAMRA is expired | No
Calibration Pure Dye TAMRA performed on | 06-01-2021
Calibration Pure Dye VIC is expired | No
Calibration Pure Dye VIC performed on | 06-01-2021
Calibration ROI is expired | No
Calibration ROI performed on | 06-01-2021
Calibration Uniformity is expired | No
Calibration Uniformity performed on | 06-01-2021
Chemistry | TAQMAN
Date Created | 2022-03-04 10:21:51 AM CST
Experiment Barcode
Experiment Comment
Experiment File Name | D:\Users\INSTR-ADMIN\Desktop\Zoe QS3 data\1110304\SMA QC test_partA2X_RNP05X_20220304_Zoe.eds
Experiment Name | SMA QC test_partA2X_RNP05X_20220304_Zoe
Experiment Run End Time | 2022-03-04 10:13:25 AM CST
Experiment Type | Standard Curve
Instrument Name | 272325672
Instrument Serial Number | 272325672
Instrument Type | QuantStudio™ 3 System
Passive Reference | ROX
Post-read Stage/Step
Pre-read Stage/Step
Quantification Cycle Method | Ct
Signal Smoothing On | true
Stage/ Cycle where Ct Analysis is performed | Stage2, Step2
User Name
Well | Well Position | Omit | Sample Name | Target Name | Task | Reporter | Quencher | CT | Ct Mean | Ct SD | Quantity | Quantity Mean | Quantity SD | Y-Intercept | R(superscript 2) | Slope | Efficiency | Automatic Ct Threshold | Ct Threshold | Automatic Baseline | Baseline Start | Baseline End | Amp Status | Comments | Cq Conf | CQCONF | THOLDFAIL | EXPFAIL | NOAMP
3 | A3 | FALSE | SMN 1-4_gDNA | RNP | UNKNOWN | TAMRA | NFQ-MGB | 24.8384 | 24.8384 | [8 empty] | TRUE | 0.04 | TRUE | 3 | 21 | Amp |  | 0.98478 | N | Y | N | N
3 | A3 | FALSE | SMN 1-4_gDNA | SMN1 | UNKNOWN | FAM | NFQ-MGB | 22.2809 | 22.2809 | [8 empty] | TRUE | 0.053028 | TRUE | 3 | 19 | Amp |  | 0.98878 | N | N | N | N
3 | A3 | FALSE | SMN 1-4_gDNA | SMN2 | UNKNOWN | VIC | NFQ-MGB | Undetermined | [9 empty] | TRUE | 0.050116 | TRUE | 3 | 34 | No Amp |  | 0 | Y | N | Y | N
15 | B3 | FALSE | SMN 2-1_gDNA | RNP | UNKNOWN | TAMRA | NFQ-MGB | 25.8381 | 25.8381 | [8 empty] | TRUE | 0.04 | TRUE | 3 | 23 | Amp |  | 0.98026 | N | Y | N | N
15 | B3 | FALSE | SMN 2-1_gDNA | SMN1 | UNKNOWN | FAM | NFQ-MGB | Undetermined | [9 empty] | TRUE | 0.053028 | TRUE | 3 | 34 | No Amp |  | 0 | Y | N | Y | Y
15 | B3 | FALSE | SMN 2-1_gDNA | SMN2 | UNKNOWN | VIC | NFQ-MGB | 24.2057 | 24.2057 | [8 empty] | TRUE | 0.050116 | TRUE | 3 | 20 | Amp |  | 0.9937 | N | N | N | N
27 | C3 | FALSE | partA-600-12_gDNA | RNP | UNKNOWN | TAMRA | NFQ-MGB | Undetermined | [9 empty] | TRUE | 0.04 | TRUE | 3 | 34 | No Amp |  | 0 | Y | Y | Y | Y
27 | C3 | FALSE | partA-600-12_gDNA | SMN1 | UNKNOWN | FAM | NFQ-MGB | 22.07 | 22.07 | [8 empty] | TRUE | 0.053028 | TRUE | 3 | 19 | Amp |  | 0.98641 | N | N | N | N
27 | C3 | FALSE | partA-600-12_gDNA | SMN2 | UNKNOWN | VIC | NFQ-MGB | 23.421 | 23.421 | [8 empty] | TRUE | 0.050116 | TRUE | 3 | 20 | Amp |  | 0.99221 | N | N | N | N
39 | D3 | FALSE | SMN 1+2_gDNA | RNP | UNKNOWN | TAMRA | NFQ-MGB | 24.2399 | 24.2399 | [8 empty] | TRUE | 0.04 | TRUE | 3 | 21 | Amp |  | 0.97204 | N | Y | N | N
39 | D3 | FALSE | SMN 1+2_gDNA | SMN1 | UNKNOWN | FAM | NFQ-MGB | 22.1496 | 22.1496 | [8 empty] | TRUE | 0.053028 | TRUE | 3 | 19 | Amp |  | 0.99198 | N | N | N | N
39 | D3 | FALSE | SMN 1+2_gDNA | SMN2 | UNKNOWN | VIC | NFQ-MGB | 24.0189 | 24.0189 | [8 empty] | TRUE | 0.050116 | TRUE | 3 | 21 | Amp |  | 0.98812 | N | N | N | N
41 | D5 | FALSE | NTC | RNP | UNKNOWN | TAMRA | NFQ-MGB | Undetermined | [9 empty] | TRUE | 0.04 | TRUE | 3 | 34 | No Amp |  | 0 | Y | Y | Y | Y
41 | D5 | FALSE | NTC | SMN1 | UNKNOWN | FAM | NFQ-MGB | Undetermined | [9 empty] | TRUE | 0.053028 | TRUE | 3 | 34 | No Amp |  | 0 | Y | N | Y | Y
41 | D5 | FALSE | NTC | SMN2 | UNKNOWN | VIC | NFQ-MGB | Undetermined | [9 empty] | TRUE | 0.050116 | TRUE | 3 | 34 | No Amp |  | 0 | Y | N | Y | Y
51 | E3 | FALSE | SMN 1+2+partA_gDNA | RNP | UNKNOWN | TAMRA | NFQ-MGB | 24.1288 | 24.1288 | [8 empty] | TRUE | 0.04 | TRUE | 3 | 21 | Amp |  | 0.98259 | N | Y | N | N
51 | E3 | FALSE | SMN 1+2+partA_gDNA | SMN1 | UNKNOWN | FAM | NFQ-MGB | 21.1303 | 21.1303 | [8 empty] | TRUE | 0.053028 | TRUE | 3 | 18 | Amp |  | 0.98494 | N | N | N | N
51 | E3 | FALSE | SMN 1+2+partA_gDNA | SMN2 | UNKNOWN | VIC | NFQ-MGB | 22.6918 | 22.6918 | [8 empty] | TRUE | 0.050116 | TRUE | 3 | 20 | Amp |  | 0.99251 | N | N | N | N
63 | F3 | FALSE | #24 | RNP | UNKNOWN | TAMRA | NFQ-MGB | 25.8853 | 25.8853 | [8 empty] | TRUE | 0.04 | TRUE | 3 | 23 | Amp |  | 0.97772 | N | Y | N | N
63 | F3 | FALSE | #24 | SMN1 | UNKNOWN | FAM | NFQ-MGB | 22.7888 | 22.7888 | [8 empty] | TRUE | 0.053028 | TRUE | 3 | 20 | Amp |  | 0.97147 | N | N | N | N
63 | F3 | FALSE | #24 | SMN2 | UNKNOWN | VIC | NFQ-MGB | 24.0813 | 24.0813 | [8 empty] | TRUE | 0.050116 | TRUE | 3 | 21 | Amp |  | 0.99007 | N | N | N | N
75 | G3 | FALSE | #25 | RNP | UNKNOWN | TAMRA | NFQ-MGB | 26.4243 | 26.4243 | [8 empty] | TRUE | 0.04 | TRUE | 3 | 24 | Amp |  | 0.96744 | N | Y | N | N
75 | G3 | FALSE | #25 | SMN1 | UNKNOWN | FAM | NFQ-MGB | 23.3184 | 23.3184 | [8 empty] | TRUE | 0.053028 | TRUE | 3 | 19 | Amp |  | 0.98826 | N | N | N | N
75 | G3 | FALSE | #25 | SMN2 | UNKNOWN | VIC | NFQ-MGB | 25.9182 | 25.9182 | [8 empty] | TRUE | 0.050116 | TRUE | 3 | 23 | Amp |  | 0.97847 | N | N | N | N
87 | H3 | FALSE | #26 | RNP | UNKNOWN | TAMRA | NFQ-MGB | 25.5506 | 25.5506 | [8 empty] | TRUE | 0.04 | TRUE | 3 | 22 | Amp |  | 0.96565 | N | Y | N | N
87 | H3 | FALSE | #26 | SMN1 | UNKNOWN | FAM | NFQ-MGB | 23.9041 | 23.9041 | [8 empty] | TRUE | 0.053028 | TRUE | 3 | 21 | Amp |  | 0.97243 | N | N | N | N
87 | H3 | FALSE | #26 | SMN2 | UNKNOWN | VIC | NFQ-MGB | 24.6428 | 24.6428 | [8 empty] | TRUE | 0.050116 | TRUE | 3 | 21 | Amp |  | 0.99006 | N | N | N | N

## Results 整理
Well | Well Position | Omit | Sample Name | Target Name | Task | Reporter | Quencher | CT | Sample Name | Dilution | Expected Molecule/ul | SMN1 | SMN2 | RNP | Delta Ct (RNP-SMN1) | Delta Ct (RNP-SMN2) | Remark |  | Sample Name | SMN1 | SMN2 | RNP | Delta Ct (RNP-SMN1) | Delta Ct (RNP-SMN2) | Remark
3 | A3 | FALSE | SMN 1-4_gDNA | RNP | UNKNOWN | TAMRA | NFQ-MGB | 24.8384 | SMN 1-4_gDNA |  |  | 22.2809 | Undetermined | 24.8384 | 2.5574 | #VALUE! | RNP-SMN1 |  | SMN 1-4_gDNA | 22.3749 | Undetermined | 25.1153 | 2.7404 | #VALUE! | RNP-SMN1
3 | A3 | FALSE | SMN 1-4_gDNA | SMN1 | UNKNOWN | FAM | NFQ-MGB | 22.2809 | SMN 2-1_gDNA |  |  | Undetermined | 24.2057 | 25.8381 | #VALUE! | 1.6323 | RNP-SMN2 |  | SMN 2-1_gDNA | Undetermined | 24.3781 | 25.4176 | #VALUE! | 1.0395 | RNP-SMN2
3 | A3 | FALSE | SMN 1-4_gDNA | SMN2 | UNKNOWN | VIC | NFQ-MGB | Undetermined | partA-600-12_gDNA |  |  | 22.07 | 23.421 | Undetermined | #VALUE! | #VALUE! | SMN1-SMN2 |  | partA-600-12_gDNA | 22.8282 | 24.8367 | Undetermined | #VALUE! | #VALUE! | SMN1-SMN2
15 | B3 | FALSE | SMN 2-1_gDNA | RNP | UNKNOWN | TAMRA | NFQ-MGB | 25.8381 | SMN 1+2_gDNA |  |  | 22.1496 | 24.0189 | 24.2399 | 2.0903 | 0.22099 | RNP-SMN1-RNP-SMN2 (1:1) |  | SMN 1+2_gDNA | 21.916 | 24.1381 | 24.2321 | 2.3161 | 0.094021 | RNP-SMN1-RNP-SMN2 (1:1)
15 | B3 | FALSE | SMN 2-1_gDNA | SMN1 | UNKNOWN | FAM | NFQ-MGB | Undetermined | SMN 1+2+partA_gDNA |  |  | 21.1303 | 22.6918 | 24.1288 | 2.9985 | 1.437 | RNP-SMN1-RNP-SMN2-SMN1-SMN2 (2:2) |  | SMN 1+2+partA_gDNA | 21.2229 | 23.4231 | 24.3218 | 3.0989 | 0.89873 | RNP-SMN1-RNP-SMN2-SMN1-SMN2 (2:2)
15 | B3 | FALSE | SMN 2-1_gDNA | SMN2 | UNKNOWN | VIC | NFQ-MGB | 24.2057 | [6 empty] | 3.9067 | 2.6531 | 3:3 | [5 empty] | 3.8817 | 1.7034 | 3:3
27 | C3 | FALSE | partA-600-12_gDNA | RNP | UNKNOWN | TAMRA | NFQ-MGB | Undetermined | [6 empty] | 4.8148 | 3.8692 | 4:4 | [5 empty] | 4.6645 | 2.5082 | 4:4
27 | C3 | FALSE | partA-600-12_gDNA | SMN1 | UNKNOWN | FAM | NFQ-MGB | 22.07 | [5 empty] | difference | 0.90817 | 1.2161 | [5 empty] | difference | 0.78282 | 0.80471
27 | C3 | FALSE | partA-600-12_gDNA | SMN2 | UNKNOWN | VIC | NFQ-MGB | 23.421
39 | D3 | FALSE | SMN 1+2_gDNA | RNP | UNKNOWN | TAMRA | NFQ-MGB | 24.2399
39 | D3 | FALSE | SMN 1+2_gDNA | SMN1 | UNKNOWN | FAM | NFQ-MGB | 22.1496 | #24_2:3 (43.2ng/ul) |  |  | 22.7888 | 24.0813 | 25.8853 | 3.0965 | 1.8041 | 2:2 | 2:3
39 | D3 | FALSE | SMN 1+2_gDNA | SMN2 | UNKNOWN | VIC | NFQ-MGB | 24.0189 | #25_2:1 (36.5ng/ul) |  |  | 23.3184 | 25.9182 | 26.4243 | 3.1059 | 0.50604 | 2:1 | 2:1 | SMN 1+2_gDNA_1/11 | 22.3764 | 23.4529 | 24.2633 | 1.8869 | 0.81044
41 | D5 | FALSE | NTC | RNP | UNKNOWN | TAMRA | NFQ-MGB | Undetermined | #26_1:2 (29.2ng/ul) |  |  | 23.9041 | 24.6428 | 25.5506 | 1.6465 | 0.90788 | 1:1 | 1:1 | SMN 1+2+partA_gDNA_1/11 | 21.2587 | 22.823 | 24.4584 | 3.1996 | 1.6354
41 | D5 | FALSE | NTC | SMN1 | UNKNOWN | FAM | NFQ-MGB | Undetermined | [14 empty] | 4.5124 | 2.4603
41 | D5 | FALSE | NTC | SMN2 | UNKNOWN | VIC | NFQ-MGB | Undetermined | [14 empty] | 5.8251 | 3.2852
51 | E3 | FALSE | SMN 1+2+partA_gDNA | RNP | UNKNOWN | TAMRA | NFQ-MGB | 24.1288 | [13 empty] | difference | 1.3127 | 0.82492
51 | E3 | FALSE | SMN 1+2+partA_gDNA | SMN1 | UNKNOWN | FAM | NFQ-MGB | 21.1303 | Sample Name | Dilution | Expected Molecule/ul | SMN1 | SMN2 | RNP
51 | E3 | FALSE | SMN 1+2+partA_gDNA | SMN2 | UNKNOWN | VIC | NFQ-MGB | 22.6918 | SMN 1-4_gDNA |  |  | 22.2809 | Undetermined | 24.8384
63 | F3 | FALSE | #24 | RNP | UNKNOWN | TAMRA | NFQ-MGB | 25.8853 | SMN 2-1_gDNA |  |  | Undetermined | 24.2057 | 25.8381
63 | F3 | FALSE | #24 | SMN1 | UNKNOWN | FAM | NFQ-MGB | 22.7888 | partA-600-12_gDNA |  |  | 22.07 | 23.421 | Undetermined
63 | F3 | FALSE | #24 | SMN2 | UNKNOWN | VIC | NFQ-MGB | 24.0813 | SMN 1+2_gDNA |  |  | 22.1496 | 24.0189 | 24.2399 | [6 empty] | Sample Name | SMN1 | SMN2 | RNP
75 | G3 | FALSE | #25 | RNP | UNKNOWN | TAMRA | NFQ-MGB | 26.4243 | SMN 1+2+partA_gDNA |  |  | 21.1303 | 22.6918 | 24.1288 | [6 empty] | SMN 1-4_gDNA | 22.3749 | Undetermined | 25.1153
75 | G3 | FALSE | #25 | SMN1 | UNKNOWN | FAM | NFQ-MGB | 23.3184 | [12 empty] | SMN 2-1_gDNA | Undetermined | 24.3781 | 25.4176
75 | G3 | FALSE | #25 | SMN2 | UNKNOWN | VIC | NFQ-MGB | 25.9182 | #24 |  |  | 22.7888 | 24.0813 | 25.8853 | [6 empty] | partA-600-12_gDNA | 22.8282 | 24.8367 | Undetermined
87 | H3 | FALSE | #26 | RNP | UNKNOWN | TAMRA | NFQ-MGB | 25.5506 | #25 |  |  | 23.3184 | 25.9182 | 26.4243 | [6 empty] | SMN 1+2_gDNA | 21.916 | 24.1381 | 24.2321
87 | H3 | FALSE | #26 | SMN1 | UNKNOWN | FAM | NFQ-MGB | 23.9041 | #26 |  |  | 23.9041 | 24.6428 | 25.5506 | [6 empty] | SMN 1+2+partA_gDNA | 21.2229 | 23.4231 | 24.3218
87 | H3 | FALSE | #26 | SMN2 | UNKNOWN | VIC | NFQ-MGB | 24.6428 | NTC |  |  | Undetermined | Undetermined | Undetermined | [6 empty] | SMN 1+2_gDNA_1/11 | 22.3764 | 23.4529 | 24.2633
[21 empty] | SMN 1+2+partA_gDNA_1/11 | 21.2587 | 22.823 | 24.4584
[21 empty] | NTC | Undetermined | Undetermined | Undetermined

## Plasmid DNA 序列稀釋
2022.03.02 下午04:30 測得的濃度
Plasmid DNA | Conc. (ng/ul) | Insert size (bp) | Total length (bp) | Remark
(Insert size+pGEM-T)
SMN1-4 | 690 | 524 | 3539 | RNP-SMN1
SMN2-1 | 858 | 524 | 3539 | RNP-SMN2
partA-600-12 | 446 | 630 | 3645 | SMN1-SMN2
236254.4269
moles dsDNA (mol) = mass of dsDNA (g)/((length of dsDNA (bp) x 617.96 g/mol/bp) + 36.04 g/mol) | [5 empty] | 181200 | 170169.2725 | 1.0648
DNA copy number = moles of dsDNA x 6.022e23 molecules/mol
*SMN1-4, SMN2-1, partA-600-12 三個的copy數需差不多
| Stock | #1 | #2 | #3 | #4 | #5 | #5-2
序列稀釋 | 1 | 1/10 | 1/100 | 1/1,000 | 1/10,000 | 1/100,000
Conc. (ng/ul) | 690 | 69 | 6.9 | 0.69 | 0.069 | 0.0069 | 0.000858
SMN1-4 (Molecule/ul) | 189994818830.2526 | 18999481883.0253 | 1899948188.3025 | 189994818.8303 | 18999481.883 | 1899948.1883 | 236254.4269 | 8.042
(每1 ul有多少Copy Number) | bar: #5-2
X ul plasmid DNA |  | 1 | 5 | 5 | 5 | 5 | 25 | 236254.4269
X ul 1xTE buffer |  | 9 | 45 | 45 | 45 | 45 | 176.049
| Stock | #1 | #2 | #3 | #4 | #5 | #6
序列稀釋 | 1 | 1/10 | 1/100 | 1/1,000 | 1/10,000 | 1/100,000 | 0.0001
Conc. (ng/ul) | 858 | 85.8 | 8.58 | 0.858 | 0.0858 | 0.00858 | 0.000858
SMN2-1 (Molecule/ul) | 236254426893.2706 | 23625442689.3271 | 2362544268.9327 | 236254426.8933 | 23625442.6893 | 2362544.2689 | 236254.4269
(每1 ul有多少Copy Number) | bar: #6
X ul plasmid DNA |  | 2 | 5 | 5 | 5 | 5 | 25
X ul 1xTE buffer |  | 18 | 45 | 45 | 45 | 45 | 225
| Stock | #1 | #2 | #3 | #4 | #5 | #5-2
序列稀釋 | 1 | 1/10 | 1/100 | 1/1,000 | 1/10,000 | 1/100,000
Conc. (ng/ul) | 446 | 44.6 | 4.46 | 0.446 | 0.0446 | 0.00446 | 0.0017674
partA-600-12 (Molecule/ul) | 119236924032.4698 | 11923692403.247 | 1192369240.3247 | 119236924.0325 | 11923692.4032 | 1192369.2403 | 472508.8538 | 2.5235
(每1 ul有多少Copy Number) | bar: #5-2
X ul plasmid DNA |  | 2 | 5 | 5 | 5 | 5 | 38 | 472508.8538
X ul 1xTE buffer |  | 18 | 45 | 45 | 45 | 45 | 57.8924
*Standard DNA + bacteria gDNA= 5ng/ul ( 25ng/5ul )。
plasmid DNA | 配製25ng/5ul | #5: X; #5-2: 將配製體積放大 100 倍
| (plasmid DNA + bacteria gDNA) | #5: bacteria gDNA濃度(ng/ul); #5-2: (plasmid DNA + bacteria gDNA)
SMN1-4 #5-2 | 1ul*SMN1-4 (ng/ul) + 4ul*X(ng/ul) | #5: 6.2498; #5-2: 5ul*SMN1-4 (ng/ul) + 20ul*X(ng/ul)
SMN2-1#6 | 1ul*SMN2-16(ng/ul) + 4ul* X(ng/ul) | #5: 6.2498; #5-2: 5ul*SMN2-1 (ng/ul) + 20ul*X(ng/ul)
partA-600-12 #5-2 | 1ul*partA-600-12 (ng/ul) + 4ul* X(ng/ul) | #5: 6.2496; #5-2: 5ul*partA-600-12 (ng/ul) + 20ul*X(ng/ul)
SMN1+ SMN2 | 1ul*SMN1-4(ng/ul) + 1ul*SMN2-1(ng/ul) + 3ul* X(ng/ul) | #5: 8.3328; #5-2: 100ul*SMN1-4 (ng/ul) + 100ul*SMN2-1 (ng/ul) + 300ul*X (ng/ul)
SMN1+SMN2+partA-600-12 | 1ul*SMN1-4(ng/ul) + 1ul*SMN2-1(ng/ul) + 1ul*partA-600-12 (ng/ul) + 2ul* X(ng/ul) | #5: 12.4983; #5-2: 100ul*SMN1-4 (ng/ul) + 100ul*SMN2-1 (ng/ul) + 100ul*partA-600-12 (ng/ul) + 200ul*X (ng/ul)
*bacteria gDNA濃度 X ng/ul 稀釋過程
bacteria gDNA stock(ng/ul): | 1080
| for SMN1、SMN2、partA-600-12 | #2: for SMN1+SMN2; #4: for SMN1+SMN2+partA-600-12 | SMN1-4 #5-2 | SMN2-1#6 | partA-600-12 #5-2
bacteria gDNA濃度(ng/ul) | 62.4979 | 6.2498 | 83.3276 | 8.3328 | 124.9826 | 12.4983 |  | 205 | 205 | 105
X ul bacteria gDNA | 2 | 7 | 3 | 35 | 4 | 25
X ul 1xTE buffer | 32.5612 | 63 | 35.8827 | 315 | 30.5648 | 225

## assay稀釋,上機條件
2022-03-04
ACCUiN SMA assay (20x)-RNP probe TAMRA |  |  |  |  | DNA | 上機DNA體積ul | 上機NFW體積ul | Reaction Mix
Primer/Probe | Stock conc. (uM) | Conc. in.rxn (uM) | 20 rxn | 1 | SMN1-4+gDNA (5ng/ul)_3/3-2 | 5 | 4 | Reagent |  | Volume (uL) per reaction |  | 9.5
SMN1/2 primer_F | 100 | 0.5 | 2 | 2 | SMN2-1+gDNA (5ng/ul)_3/3-2 | 5 | 4 | 2X qPCRBIO Probe Mix + ROX II |  | 10 |  | 95
SMN1/2 primer_R | 100 | 0.5 | 2 | 3 | partA-600-12+gDNA (5ng/ul)_3/3-2 | 5 | 4 | Assay (20x) |  | 1 |  | 9.5
SMN1 probe | 100 | 0.4 | 1.6 | 4 | SMN1-4+SMN2-1+gDNA (5ng/ul)_3/3-2 | 5 | 4 | Plasmid DNA |  | 參考左表
SMN2 probe | 100 | 0.2 | 0.8 | 5 | SMN1-4+SMN2-1+partA-600-12+gDNA (5ng/ul)_3/3-2 | 5 | 4 | Nuclease-Free Water |  | 參考左表
RNP primer_F | 100 | 0.25 | 1 | 6 | #24_2:3 (43.2ng/ul) | 1.16 | 7.84 | Total |  | 20
RNP primer_R | 100 | 0.25 | 1 | 7 | #25_2:1 (36.5ng/ul) | 1.37 | 7.63
RNP probe_TAMRA | 100 | 0.1 | 0.4 | 8 | #26_1:2 (29.2ng/ul) | 1.71 | 7.29
H2O | - | - | 11.2 | 9 | NTC | 0 | 9
qPCR program
[6 empty] | SNN1: FAM |  | Stage | Step | Temperature (°C) | Time
[6 empty] | SNN2: VIC |  | Hold | Activation | 95 | 3 min
[6 empty] | RNP: TAMRA |  | Cycling*35 | Denaturation | 95 | 5 sec
[6 empty] | Refrencen: ROX |  |  | Extension (capture) | 60 | 30 sec